Quality of Prostate MRI: Is the PI-RADS Standard Sufficient?

Jonathan Sackett B.S.1,2; Joanna H. Shih PhD1; Sarah E. Reese PhD3; Jeffrey R. Brender PhD1; Stephanie Harmon PhD1,4; Tristan Barrett MD5; Mehmet Coskun MD6; Manuel Madariaga MD7; Jamie Marko MD1; Yan Mee Law MD8; Evrim B. Turkbey MD1; Sherif Mehralivand MD1; Thomas Sanford MD1; Nathan Lay PhD1; Peter A. Pinto MD1; Bradford J. Wood MD1; Peter L. Choyke MD1; Baris Turkbey MD1

1 National Cancer Institute, National Institutes of Health, Bethesda, MD, USA

2 Loma Linda University School of Medicine, Loma Linda, CA, USA

3 General Dynamics Information Technology, Falls Church, VA, USA

4 Leidos Biomedical Research, Inc., NCI Campus at Frederick, Clinical Research Directorate/Clinical Monitoring Research Program, Bethesda, MD, USA

5 University of Cambridge School of Clinical Medicine, Cambridge UK

6 Department of Radiology, Dr. Behcet Uz Child Disease and Pediatric Surgery Training and Research Hospital, University of Health Sciences, İzmir, Turkey

7 Department of Radiology, San Jose Hospital, Santiago, Chile

8 Department of Diagnostic Radiology, Singapore General Hospital, Singapore.

Corresponding Author:

Baris Turkbey MD

Abbreviations:

PCa – Prostate Cancer, CSPCa – Clinically significant prostate cancer, TRUS – Transrectal ultrasound, mpMRI – multi-parametric MRI, PI-RADS - Prostate Imaging Reporting and Data System, MTS – Minimum Technical Standards, FOV – field-of-view, IPRP - in-plane resolution phase, IPRF - in-plane resolution frequency, ST – Slice Thickness, SG – Slice Gap

Acknowledgements: The content of this publication does not necessarily reflect the views or policies of the Department of Health and Human Services, nor does mention of trade names, commercial products, or organizations imply endorsement by the US Government.

Grant Support: This research was made possible through the NIH Medical Research Scholars Program, a public-private partnership supported jointly by the NIH and contributions to the Foundation for the NIH from the Doris Duke Charitable Foundation (DDCF Grant #2014194), the American Association for Dental Research, the Colgate-Palmolive Company, Genentech, Elsevier, and other private donors.

Running Title: Quality of Prostate MRI: Is the PI-RADS Standard Sufficient?

Quality of Prostate MRI: Is the PI-RADS Standard Sufficient?

Abstract

Rationale and Objective: The Prostate Imaging Reporting and Data System version 2 (PI-RADSv2) published a set of Minimum Technical Standards (MTS) to improve image quality and reduce variability in Multi-parametric prostate MRI (mpMRI). The effect of PIRADSv2 MTS on image quality has not been validated. We aimed to determine whether adherence to PI-RADSv2 MTS improves study adequacy and perceived quality.

Methods: 62 prostate MRI examinations including T2W and DWI consecutively referred to our center from 62 different institutions within a 12-month period (September 2017-September 2018) were included. 6 readers assessed images as adequate or inadequate for use in PCa detection and a numerical image quality ranking was given using a 1-5 scale. The PI-RADSv2 Minimum Technical Standards were synthesized into sets of 7 and 10 rules for T2W and DWI, respectively. Image adherence was assessed using DICOM metadata. Statistical analysis of survey results and image adherence was performed based on reader quality scoring (Kendall Rank Tau-b) and reader adequate scoring (Wilcoxon test for association) for T2 and DWI quality assessment.

Results: Out of 62 images, 52 (83%) T2W and 38 (61%) DW images were rated to be adequate by a majority of readers. Reader adequacy scores showed no significant association with adherence to PI-RADSv2. There was a weak (tau-b=0.22) but significant (p-value=0.01) correlation between adherence to PIRADSv2 MTS and image quality for T2W. Studies following all PI-RADSv2 T2W rules achieved a higher median average quality score (3.58 for 7/7 vs. 3.0 for <7/7, p=0.012). No statistical relationship with PI-RADSv2 MTS adherence and DWI quality was found.

Conclusion: Among 62 sites performing prostate MRI, few were considered of high quality, but the majority were considered adequate. DWI showed considerably lower rates of adequate studies in the sample. Adherence to PI-RADSv2 Minimum Technical Standards did not increase the likelihood of having a qualitatively adequate T2W or DWI.

Keywords: Prostate, Quality Control, MRI, Multi-parametric, PI-RADS, Diffusion Weighted Imaging

Abbreviation Key:

PCa – Prostate Cancer, CSPCa – Clinically significant prostate cancer, TRUS – Transrectal ultrasound, mpMRI – multi-parametric MRI, PI-RADS – Prostate Imaging Reporting and Data System, MTS – Minimum Technical Standards, FOV – field-of-view, IPRP - in-plane resolution phase, IPRF - in-plane resolution frequency, ST – Slice Thickness, SG – Slice Gap

Introduction

Approximately, 175,000 men will be diagnosed with prostate cancer (PCa) in the United States in 2019 making it the most common solid organ malignancy in men. However, only approximately 31,600 men will die from PCa in the same period.(1) Trying to distinguish clinically significant PCa (CSPCa) from indolent PCa is a central challenge in prostate cancer management. Traditional screening methods which rely on prostate-specific antigen (PSA) screening and transrectal ultrasound (TRUS) guided biopsy results in over-diagnosis and overtreatment.(2) The introduction of prostate multi-parametric MRI (mpMRI) improves the detection of clinically significant cancers while reducing the diagnosis of indolent cancer. The creation of the Prostate Imaging Reporting and Data System Version 1 (PI-RADSv1) was an effort to standardize the interpretation of mpMRI. PI-RADSv2 simplified the original system, reducing the role of dynamic contrast enhanced MRI and creating a single composite score for each lesion. Importantly, PI-RADSv2 also suggested standards for the acquisition of the individual MRI sequences based on the expert opinion of the PI-RADSv2 committee members.(3, 4)

PI-RADSv2 has consistently been shown to have a high negative predictive value allowing physicians to triage patients and avoid biopsies in patients unlikely to have prostate cancer.(5, 6) Similarly, other studies have shown that PI-RADSv2 in the context of MRI guided biopsies has increased sensitivity for CSPCa over traditional TRUS guided biopsy methods.(7–10)

However, the problem of variable image quality among centers is a major limitation of mpMRI. Low quality prostate MRIs could easily limit lesion localization lowering diagnostic yield in MRI guided biopsies resulting in no patient benefit over standard ultrasound guidance. As mpMRI disseminates to a wider user-base, scans that are of limited diagnostic value may become more common. Image quality is notoriously difficult to assess but clearly could impact the results of recent large studies that showed little difference between MRI-guided and TRUS-guided biopsies.(6, 11)

The PI-RADSv2 Minimum Technical Standards (MTS) were created with the goal of standardizing imaging protocols and reducing image quality variance.(4) However, Esses et al. found that adherence to the guidelines varied widely across different institutions with some standards only achieving 17% adherence.(12) However, it is reasonable to ask whether adherence to the standards solves the quality problem. It is possible that the sites not adhering to the standards may be doing so because they find that the MTS do not work well in their setting.

Therefore, this study was designed to evaluate variability in quality of prostate mpMRI and relate quality to the degree of adherence to the PI-RADSv2 MTS for T2 and DWI sequences, the two major components of mpMRI.

Materials and Methods

Study Population

We identified a cross-sectional cohort of 62 prostate MRI examinations performed at outside imaging facilities which had been consecutively referred to our center for secondary interpretation within a 12-month period (September 2017-September 2018) including both T2 weighted (T2W) and diffusion weighted image (DWI) sequences. All patients enrolled at the National Cancer Institute Center for Cancer Research under an Institutional Review Board approved trial (National Clinical Trials identifier NCT02594202) and signed an informed consent form. After anonymization, DICOM files for T2W and raw DWI sequences were uploaded to the local PACS system. The age of patients ranged from 48 to 80 with a mean age of 66 (SD ±8). The distribution of MRI vendors for acquisition was Siemens 58%, General Electric (GE) 35%, and Philips 6%. 90% of exams were performed using a 3 Tesla scanner (Table 1).

Assessment of PIRADSv2 MTS Adherence

DICOM files and metadata were used to assess adherence to PI-RADSv2 MTS for T2W and DWI sequences. This assessment was aided by an in-house developed automated assessment tool which assessed adherence to PI-RADSv2 MTS. In addition to the PI-RADSv2 MTS, we recorded the phase encoding direction for each image as we hypothesized that the phase encoding direction might have an effect on image quality and severity of distortion on DWI sequences with horizontal phase encoding direction (right–>left) performing better than vertical phase encoding direction (anterior –> posterior).

Survey Design

Three experienced (TB, JM, YML) and three novice readers (MC, ET, MM) from 5 different institutions (non-contributors to the test cases) were recruited to assess each image and sequence on an array of subjective image quality characteristics. Readers assessed images as either “adequate” or “inadequate” for use in prostate cancer detection. Readers’ assessment of image quality was ranked on a 1-5 Likert scale with 5 being the best and 1 being impossible to interpret. Readers were trained on a different set of reference images before being given the test set. For training one expert radiologist (BT) selected cases considered to represent the spectrum of image quality (1-4). Since scores of “5” were unusual, readers were instructed to give a score of 5 when the image quality exceeded the reference image for the quality score of 4. Readers also assessed images according to visual characteristics and for the presence and severity of perceived artifacts. The visual characteristics chosen were blurriness and contrast.(13, 14) The perceived artifacts chosen were noise, motion, geometric distortion, and aliasing/ghosting. The visual characteristics and perceived artifacts were graded on a 1-5 scale with a score of 1 indicating that the specific artifact was so severe that the image was impossible to interpret and 5 indicating the absence of the artifact. Readers were not told whether an image was compliant with the PI-RADSv2 MTS. However, readers were aware of the number of b-values obtained and the b-value of the highest b-value image.

Statistical Analysis

Association between PI-RADSv2 adherence, image adequacy, and image quality were assessed by Kendall's tau-b rank correlation, Wilcoxon rank sum test, or chi-square test accounting for inter-reader correlation arising from multiple readers reviewing the same cases.(15, 16) Reader agreement on image quality was determined by Fleiss' kappa. Standard errors and 95% confidence intervals of kappa statistics were estimated from 2000 bootstrap samples by random sampling on the case level. P-values <0.05 were considered statistically significant.

Results

Adequacy and Quality of prostate mpMRI

Of the 62 studies, 52 (84%) T2W and 38 (61%) DWI were rated to be diagnostically adequate. However, for only 35 (56%) of the studies, both sequences were rated as adequate by a majority. For T2W, only 10 (16%) were scored as high quality (score >3) by a majority of readers. For DWI, only 6 (9%) were scored as high quality by a majority of readers.

Reader Agreement

Reader agreement on image adequacy was fair for T2W (=0.40) and DWI (=0.39). Reader agreement on image quality was poor in T2W (=0.17) and fair in DWI (=0.21). Reader agreement on image blurriness and contrast was poor in both T2 and DWI (Table 2) suggesting there is considerable disagreement regarding image quality.

Adherence Rates

Adherence to PI-RADSv2 MTS varied widely. Eighteen studies met all 7 T2 standards and 6 studies met all 10 DWI standards. The T2W imaging standards with the lowest rates of adherence were in-plane resolution phase ≤ 0.4mm and no interslice gap, with adherence rates of 52% and 61% respectively. For DWI, the standards with the lowest rates of adherence were field-of-view (FOV) ≤ 220mm and high-b value ≥ 1400s/mm2 with adherence rates of 35% and 50% respectively (Table 3).

Adherence and Diagnostic Adequacy

There was no significant association between adherence to PI-RADSv2 MTS and image adequacy for either T2 (p-value=0.227) or DWI (p-value=0.304) indicating that following the technical standards had no effect on the likelihood of an image being scored as diagnostically adequate (Figure 1, Table 4).

Adherence and Quality

There was a significant association between adherence to PI-RADSv2 technical standards and image quality for T2W (p-value=0.010) but not for DWI (p-value=0.943) indicating that T2W images were more likely to receive a higher quality score if they followed the PIRADSv2 MTS while DWI images were not. However, the strength of the correlation (as determined by the Tau-b) between T2 image quality and adherence was only 0.22 (SE 0.08), indicating a weak correlation (Figure 2, Table 4). Significant associations were also found between the perceived quality features and perceived artifact metrics of Noise, Distortion, Aliasing, Blurriness, and Contrast for T2 images (Table 4).

For DWI, not only was there no significant association between adherence to PI-RADSv2 technical standards and image quality, but the tau-b degree of correlation between the two measures was -0.01 (SE 0.08) signifying a near complete absence of correlation (Table 4). No significant associations were found for any of the DWI quality/artifact metrics and adherence (Figure 3, Table 4).

Performance of Individual PIRADSv2 Standards

There was no significant association between any specific PI-RADSv2 minimum technical standard and image adequacy or image quality (Table 5).

Individual PI-RADSv2 standards were also analyzed for their effect on image distortion. There was a significant association in T2W imaging with adherence to the FOV UL ≤ 200mm standard (p-value=0.003). In DWI, adherence to the TR ≥ 3000ms standard (p-value=0.033) was associated with less distorted images. There was no significant association between adherence to any other specific standard and distortion for either T2W or DWI (Table 5).

Analysis of the use of horizontal phase encoding direction found a significant association with T2W image quality (p-value=0.032) where images acquired with a horizontal phase encoding direction performed better than those with a vertical phase encoding direction. This same association also existed for distortion in T2W sequences (p-value=0.004) where images were less likely to be distorted if phase encoding direction was horizontal. However, this association did not exist for DWI (Table 5).

Perceived Artifact and Impact on Diagnostic Adequacy and Image Quality

All perceived artifacts (noise, motion, geometric distortion, and aliasing) had a significant association (p-values ≤ 0.01) for both image adequacy and quality. For T2W images, the perceived artifact with the greatest effect on image quality was noise. The tau-b correlation coefficient for noise and image quality was 0.58 (SE 0.09). For DWI, both noise and distortion had particularly strong correlations with image quality with tau-b correlation coefficients of 0.61 (SE 0.09) and 0.54 (SE 0.09) respectively (Table 6).

The frequency with which readers scored perceived artifacts (noise, motion, geometric distortion, and aliasing) as making image interpretation difficult or impossible (reader score < 3) was also assessed for both T2 and DWI. For T2, the frequency at which readers gave scores < 3 (interpretation difficult or impossible) was 19%, 13%, 6%, and 8% for noise, motion, geometric distortion, and aliasing respectively. For DWI, 30%, 10%, 27%, and 26% of reader scores were < 3 (interpretation difficult or impossible) for noise, motion, geometric distortion, and aliasing respectively.

Discussion

Performance of PIRADSv2 Minimum Technical Standards

This study shows a disheartening rate of inadequate exams among mpMRI studies of the prostate. Adherence to PI-RADSv2 MTS did not appear to play a role in improving image quality.  Many centers found a way to produce diagnostic images while essentially overlooking the PI-RADSv2 standards while many who demonstrated strict adherence to the standards had poor results. While our study findings indicated that T2W imaging adherence to the PI-RADSv2 MTS did increase the likelihood of being scored as higher quality, the actual effect was very slight. Furthermore, adherence to the technical standards did not significantly increase the probability that an image would be seen as diagnostically adequate. For DWI, the results were even more striking with adherence to the PI-RADSv2 MTS having no significant correlation with either adequacy or quality scores.

The analysis of each PIRADSv2 MTS found no significant relationship with quality or adequacy for any current individual standard. This suggests a “recipe-like” standard may not be adequate, but rather a standard based on objective imaging characteristics may be needed. Despite this, we did find a significant improvement in quality and distortion for T2W sequences using horizontal phase encoding direction suggesting this feature should be included in future PIRADS versions.

Variable Quality and Adequacy of prostate imaging

As expected, many centers struggle to achieve adequate image quality for prostate MRI. This problem is particularly acute for DWI sequences with only 61% (as opposed to 83% for T2) scored as adequate by a majority of our study readers. Additionally, among those images which are adequate for use relatively few were considered by readers to be high quality. The wide variability in image quality indicates a need for quality assurance and protocols to standardize image quality in prostate mpMRI. We suggest greater involvement of MRI vendors and MRI physicists in the improvement of mpMRI. This has long been discussed in the prostate mpMRI community.(4, 17) However, the problem has never been clearly defined.

Adherence to PIRADSv2 Minimum Technical Standards

Several of the PI-RADS standards had particularly low adherence rates suggesting they may be difficult to implement on some scanners. In T2W sequences, the in-plane resolution frequency ≤ 0.4mm had an adherence rate of only 52% and, in DWI, only 35% of images analyzed were adherent to upper limit for FOV≤220mm. This is the same pattern of adherence rates found in the previous study to assess rates of adherence to the PIRADSv2 Technical Standards by Esses et al.(12) Our results from 2018 only found improvements in adherence for resolution and slice thickness. There has been no improvement for the other standards since the 2015-2016 data collection by Esses et al.

Notably, there were T2W images that followed only a few of the standards and were still seen as adequate by all 6 readers and scored as high quality. The presence of non-adherent high-quality images indicates that the current T2 standards do not represent a true minimum standard. However, the problem of non-adherent high-quality outliers is more pronounced in the DWI dataset with many non-adherent images achieving high quality and very few adherent images achieving high quality.

Future Directions

Looking forward, it is apparent the general quality of prostate mpMRI should be improved to deliver better diagnosis in PCa care. Although providing formal recommendations for quality prostate MRI is beyond the scope of this paper, there are several observations that may be useful in improving mpMRI. Image noise was most frequently implicated in making interpretation difficult for both T2W and DWI sequences. Image noise is particularly important at high b-values in DWI. Noise is likely to become a greater factor if distortion corrections are employed as these methods typically come at a cost of decreased SNR either intrinsically or from decreased signal averaging to accommodate additional scans. Techniques to reduce noise during post-processing should also be considered.(18)

Image distortion particularly for DWI was the second most frequently implicated perceived artifact. One of the most prominent sources of distortion in prostate imaging is the inhomogeneous B0 field in the abdominal area.(19, 20) Most biological tissues are slightly diamagnetic with the exception of gas filled cavities. If the rectum is filled with gas during a scan, the difference in effective B0 fields at the rectum/prostate interface creates a magnetic field gradient which shifts intensity from one location to another creating geometric distortion with signal loss near the interface due to spin dephasing.(21) While recognized in the literature, corrections for susceptibility artifacts are not yet incorporated into the PIRADS standard.

Susceptibility artifact is important during DWI acquisition due to ubiquitous use of the echo planar imaging (EPI) sequence which magnifies magnetic inhomogeneity.(22) Susceptibility artifact can be reduced by eliminating rectal gas by bowel preparation,(23) or by the administration of a spasmolytic agent (e.g. buscopan).(24) During acquisition, we have shown that the effect of the susceptibility artifact can be mitigated by altering the phase encoding direction from vertical (anterior-posterior) to horizontal (left-right). Although not seen in our cohort, flipping the phase encoding direction can theoretically move the distortion to the lateral axis and away from the prostate. Limiting spin evolution during readout by reducing echo times,(19) restricting the FOV,(25) or by parallel imaging has the potential to diminish susceptibility effects. Using B0 field maps to adjust the signal phase in postprocessing or combining scans of opposite phase encoding polarity to cancel improper phase evolution have also been shown to be effective in minimizing distortion artifacts.(26) These and other candidate strategies need further exploration and investigation in larger studies.

Limitations

This study has several limitations. First, we evaluated the impact of adherence to each standard in a binary fashion (compliant or non-compliant) on adequacy and quality; however, we did not evaluate the effect of the actual values of image acquisition settings on quality. Second, we approached the quality evaluation for image acquisition parameters in a binary fashion using strict adherence criteria of MTS in the PI-RADS documents instead of quantifying the degree of difference from the MTS defined in PI-RADS.  We agree that this may underestimate the importance of PI-RADS MTS.  However, there is an inherent value judgement regarding how closely an MTS standard was followed and to avoid this subjectivity we opted for a more absolute criterion of compliance. Third, readers evaluated the adequacy of images for use in diagnosis, but we did not assess the actual effect on diagnostic accuracy. To do both of the above would require a vast sample size and a prospective design to avoid ascertainment bias. Moreover, we did not evaluate the quality of ADC maps directly. Since PI-RADSv2 has not specified standard methodology for the post-processing calculation of ADC maps, we decided to evaluate only image sequences that could be acquired without post-processing. Additionally, we did not get a chance to evaluate the experience level of the technologists who conducted the image acquisition since it was not possible to reach out all the 62 different centers. Finally, reader agreement on image quality evaluated by Fleiss’ kappa score was only slight or fair indicating a lack of consensus on image quality, highlighting a need for more objective quality control.

Conclusion

In conclusion, the results indicate image quality in prostate mpMRI is highly variable. Adherence to PI-RADSv2 Minimum Technical Standards has a minimal effect in reducing this variability for T2 imaging sequences and no effect for DWI sequence. Noise and distortion are the most obvious sources of problems in most prostate mpMRIs. These results suggest that simply prescribing a set of sequence specifications is unlikely to solve the problem of image quality. Image quality also depends on patient factors such as the presence of rectal gas, total-hip replacement, amount of body fat, post-biopsy hemorrhage, and abstinence from ejaculation.(27–30). It may be difficult to devise new purely technical standards to account for these factors. However, there are other quality assurance methods that can be devised such as standards for patient factors or perhaps methods which use machine learning techniques such as texture analysis or convolutional neural networks.(31, 32)

Image quality remains a significant barrier to improving prostate multi-parametric MRI. Meanwhile, scan settings should be individually tailored to each scanner and institution with PI-RADSv2 minimum technical standards used as a possible starting point for optimization and not boundaries for optimization. There is an urgent need for new quality assurance tools, involvement of MRI manufacturers, the MR physics community and individual efforts on the part of radiologists to improve mpMRI.

References:

1. Siegel RL, Miller KD, Jemal A: Cancer statistics, 2019. CA Cancer J Clin 2019; 69:7–34.

2. Shen X, Kumar P: Trade-Off between Treatment of Early Prostate Cancer and Incidence of Advanced Prostate Cancer in the Prostate Screening Era. J Urol 2016; 195:1397–1402.

3. Weinreb JC, Barentsz JO, Choyke PL, et al.: PI-RADS Prostate Imaging – Reporting and Data System: 2015, Version 2. Eur Urol 2016; 69:16–40.

4. Padhani AR, Weinreb J, Rosenkrantz AB, Villeirs G, Turkbey B, Barentsz J: Prostate Imaging-Reporting and Data System Steering Committee: PI-RADS v2 Status Update and Future Directions. Eur Urol 2019; 75:385–396.

5. Liu C, Liu S-L, Wang Z-X, et al.: Using the prostate imaging reporting and data system version 2 (PI-RIDS v2) to detect prostate cancer can prevent unnecessary biopsies and invasive treatment. Asian J Androl 2018; 20:459.

6. van der Leest M, Cornel E, Israël B, et al.: Head-to-head Comparison of Transrectal Ultrasound-guided Prostate Biopsy Versus Multiparametric Prostate Resonance Imaging with Subsequent Magnetic Resonance-guided Biopsy in Biopsy-naïve Men with Elevated Prostate-specific Antigen: A Large Prospective Mu. Eur Urol 2019; 75:570–578.

7. Mehralivand S, Bednarova S, Shih JH, et al.: Prospective Evaluation of PI-RADS Version 2 Using the International Society of Urological Pathology Prostate Cancer Grade Group System. J Urol 2017; 198:583–590.

8. Kim SH, Choi MS, Kim MJ, Kim YH, Cho SH: Validation of Prostate Imaging Reporting and Data System Version 2 Using an MRI-Ultrasound Fusion Biopsy in Prostate Cancer Diagnosis. Am J Roentgenol 2017; 209:800–805.

9. Ahmed HU, El-Shater Bosaily A, Brown LC, et al.: Diagnostic accuracy of multi-parametric MRI and TRUS biopsy in prostate cancer (PROMIS): a paired validating confirmatory study. Lancet 2017; 389:815–822.

10. Kasivisvanathan V, Rannikko AS, Borghi M, et al.: MRI-Targeted or Standard Biopsy for Prostate-Cancer Diagnosis. N Engl J Med 2018; 378:1767–1777.

11. Rouvière O, Puech P, Renard-Penna R, et al.: Use of prostate systematic and targeted biopsy on the basis of multiparametric MRI in biopsy-naive patients (MRI-FIRST): a prospective, multicentre, paired diagnostic study. Lancet Oncol 2019; 20:100–109.

12. Esses SJ, Taneja SS, Rosenkrantz AB: Imaging Facilities’ Adherence to PI-RADS v2 Minimum Technical Standards for the Performance of Prostate MRI. Acad Radiol 2018; 25:188–195.

13. Ullrich T, Quentin M, Oelers C, et al.: Magnetic resonance imaging of the prostate at 1.5 versus 3.0 T: A prospective comparison study of image quality. Eur J Radiol 2017; 90:192–197.

14. Loening AM, Litwiller D V, Saranathan M, Vasanawala SS: Increased Speed and Image Quality for Pelvic Single-Shot Fast Spin-Echo Imaging with Variable Refocusing Flip Angles and Full-Fourier Acquisition. Radiology 2017; 282:561–568.

15. Rosner B, Glynn RJ, Ting Lee M: Incorporation of Clustering Effects for the Wilcoxon Rank Sum Test: A Large‐Sample Approach. Biometrics 2003; 59:1089–1098.

16. Shih JH, Fay MP: Pearson’s chi-square test and rank correlation inferences for clustered data. Biometrics 2017; 73:822–834.

17. Rosenkrantz AB, Oto A, Turkbey B, Westphalen AC: Prostate Imaging Reporting and Data System (PI-RADS), Version 2: A Critical Look. AJR Am J Roentgenol 2016; 206:1179–83.

18. Zhang X, Peng J, Xu M, et al.: Denoise diffusion-weighted images using higher-order singular value decomposition. Neuroimage 2017; 156(January 2016):128–145.

19. Nketiah G, Selnæs KM, Sandsmark E, et al.: Geometric distortion correction in prostate diffusion-weighted MRI and its effect on quantitative apparent diffusion coefficient analysis. Magn Reson Med 2018; 79:2524–2532.

20. Rakow-Penner RA, White NS, Margolis DJA, et al.: Prostate diffusion imaging with distortion correction. Magn Reson Imaging 2015; 33:1178–1181.

21. Port JD, Pomper MG: Quantification and Minimization of Magnetic Susceptibility Artifacts on GRE Images. J Comput Assist Tomogr 2000; 24:958–964.

22. Tsougos I: Artifacts and Pitfalls in Diffusion MRI. Ser Med Phys Biomed 2018:29–54.

23. Caglic I, Hansen NL, Slough RA, Patterson AJ, Barrett T: Evaluating the effect of rectal distension on prostate multiparametric MRI image quality. Eur J Radiol 2017; 90:174–180.

24. Ullrich T, Quentin M, Schmaltz AK, et al.: Hyoscine butylbromide significantly decreases motion artefacts and allows better delineation of anatomic structures in mp-MRI of the prostate. Eur Radiol 2018; 28:17–23.

25. Korn N, Kurhanewicz J, Banerjee S, Starobinets O, Saritas E, Noworolski S: Reduced-FOV excitation decreases susceptibility artifact in diffusion-weighted MRI with endorectal coil for prostate cancer detection. Magn Reson Imaging 2015; 33:56–62.

26. Hong X, To XV, Teh I, Soh JR, Chuang KH: Evaluation of EPI distortion correction methods for quantitative MRI of the brain at high magnetic field. Magn Reson Imaging 2015; 33:1098–1105.

27. Sharif-Afshar A-R, Feng T, Koopman S, et al.: Impact of post prostate biopsy hemorrhage on multiparametric magnetic resonance imaging. Can J Urol 2015; 22:7698–702.

28. Medved M, Sammet S, Yousuf A, Oto A: MR Imaging of the Prostate and Adjacent Anatomic Structures before, during, and after Ejaculation: Qualitative and Quantitative Evaluation. Radiology 2014; 271:452–460.

29. Rosenkrantz AB, Kopec M, Kong X, et al.: Prostate cancer vs. post-biopsy hemorrhage: Diagnosis with T2- and diffusion-weighted imaging. J Magn Reson Imaging 2010; 31:1387–1394.

30. Kabakus IM, Borofsky S, Mertan F V., et al.: Does Abstinence From Ejaculation Before Prostate MRI Improve Evaluation of the Seminal Vesicles? Am J Roentgenol 2016; 207:1205–1209.

31. Esses SJ, Lu X, Zhao T, et al.: Automated image quality evaluation of T 2 -weighted liver MRI utilizing deep learning architecture. J Magn Reson Imaging 2018; 47:723–728.

32. Martirosian P, Liebgott A, Yang B, et al.: A machine-learning framework for automatic reference-free quality assessment in MRI. Magn Reson Imaging 2018; 53(July):134–147.

Figure Legends

Figure 1: The distribution of T2 (left) and DWI (right) imaging studies organized by the number of adequate votes received and by the number of PIRADSv2 Minimum Technical Standards (MTS) followed. The numbers within the grid indicate the number of studies following the number of MTS indicated (column) and received the number of votes indicated (row).

Figure 2: Relationship between adherence to T2W PI-RADSv2 Minimum Technical Standards (MTS) and average reader quality score. Select outlier images have been given color coded labels containing their average reader quality score. Green labels: Non-compliant but high quality images. Orange labels: Compliant but low quality images.

Figure 3: Relationship between adherence to DWI PI-RADSv2 Minimum Technical Standards (MTS) and average reader quality score. Select outlier images have been given color coded labels containing their average reader quality score. Green labels: Non-compliant but high quality images. Orange labels: Compliant but low quality images.

Figure 4: Examples which highlight the gaps of the PIRADSv2 Minimum Technical Standards (MTS) all acquired at 3 Teslas. (Left Column) An example of an image that is non-adherent to the PIRADSv2 MTS yet still achieved a high-quality image with a clearly visible and easily defined lesion.  (Middle Column) 100% adherent to the PIRADSv2 MTS for T2 sequences yet the T2 sequence appears noisy and the separation between the peripheral zone and transition zone cannot be appreciated. (Right Column) 100% adherent to the PIRADSv2 MTS for DWI sequences yet the low b-value sequence has significant distortion and the prostate is difficult to visualize in the high b-value sequence.

Tables

| T2 Parameters | Mean ± SD | Median | Range |
| --- | --- | --- | --- |
| Slice Thickness (ST) | 3.0 ± 0.5 | 3 | 1.1–5 |
| Field of View (FOV) | 187 ± 25 | 180 | 120–250 |
| in-plane Resolution (Phase) | 0.5 ± 0.2 | 0.4 | 0.2–1.1 |
| in-plane Resolution (Frequency) | 0.5 ± 0.2 | 0.4 | 0.2–1.1 |
| Slice Gap (SG) | Mean ± SD | No. w/ gap | Mean Gap* |
|  | 0.2 ± 0.3 | 24 | 0.60 |
| Phase Encoding Direction | # Vertical | # Horizontal | - |
|  | 15 | 47 | - |
| DWI Parameters | Mean ± SD | Median | Range |
| High B Value | 1222 ± 427 | 1400 | 500–2000 |
| # of B Values | 2.7 ± 1.0 | 2 | 2–7 |
| TE | 80.4 ± 15.6 | 83 | 47–110 |
| TR | 6438 ± 2637 | 5622 | 2800–13500 |
| Slice Thickness (ST) | 3.4 ± 0.6 | 3.25 | 2.7–6.0 |
| Field of View (FOV) | 232 ± 63 | 250 | 90–380 |
| in-plane Resolution (Phase) | 1.4 ± 0.5 | 1.3 | 0.5–2.6 |
| in-plane Resolution (Frequency) | 1.4 ± 0.5 | 1.3 | 0.5–2.6 |
| Slice Gap (SG) | Mean ± SD | No. w/ gap | Mean Gap* |
|  | 0.1 ± 0.3 | 13 | 0.58 |
| Phase Encoding Direction | # Vertical | # Horizontal | - |
|  | 47 | 15 | - |
| Machine Attributes |  |  |  |
| Magnet Strength | # of 1.5T | # of 3T | - |
|  | 6 | 56 | - |
| Vendor | GE | Siemens | Philips |
|  | 22 | 36 | 4 |

*Mean gap calculation excludes all sequences with no gap

Table 1: Summary of the range of prostate MRI acquisition parameters in this study population

|  | Agreement Metric | Adequate | Quality | Blurriness | Contrast |
| --- | --- | --- | --- | --- | --- |
| T2 | Complete Agreement* | 58% | 2% | 0% | 3% |
|  | Majority Agreement* | 84% | 47% | 50% | 50% |
|  | Fleiss’ Kappa (SE) | 0.40 (0.08) | 0.17 (0.03) | 0.17 (0.03) | 0.19 (0.03) |
|  | Conclusion† | Fair | Slight | Slight | Slight |
| DWI | Complete Agreement* | 29% | 5% | 0% | 0% |
|  | Majority Agreement* | 61% | 47% | 50% | 39% |
|  | Fleiss’ Kappa (SE) | 0.39 (0.06) | 0.21 (0.04) | 0.18 (0.03) | 0.14 (0.02) |
|  | Conclusion† | Fair | Fair | Slight | Slight |

Table 2: Reader agreement on Subjective Image Quality Survey

*Complete Agreement defined as % of cases where all readers agreed on score, Majority agreement defined as % of cases where the majority of readers agreed on score.

†Conclusion determined by Kappa score: Moderate ≥ 0.41, Fair ≥ 0.21, and Slight ≤ 0.20

| T2W Technical Standards | % Adherence |  |
| --- | --- | --- |
| magnet strength ≥ 1.5T |  | 100% |
| slice thickness (ST) ≤ 3mm |  | 92% |
| no inter-slice gap |  | 61% |
| field of view lower limit (FOV LL) ≥ 120mm |  | 100% |
| field of view upper limit (FOV UL) ≤ 200mm |  | 82% |
| in-plane resolution phase (IPRP) ≤ 0.7mm |  | 92% |
| in-plane resolution frequency (IPRF) ≤ 0.4mm |  | 52% |
| DWI Technical Standards | % Adherence |  |
| magnet strength ≥ 1.5T |  | 100% |
| High b-value images ≥1400s/mm2 |  | 50% |
| TR ≥ 3000ms |  | 94% |
| TE ≤ 90ms |  | 73% |
| slice thickness ≤ 4mm |  | 97% |
| no inter-slice gap |  | 79% |
| field of view lower limit (FOV LL) ≥ 160mm |  | 87% |
| field of view upper limit (FOV UL) ≤ 220 mm |  | 35% |
| in-plane resolution phase (IPRP) ≤ 2.5mm |  | 98% |
| in-plane resolution frequency (IPRF) ≤ 2.5mm |  | 98% |

Table 3. Summary of overall rates of adherence to PIRADSv2 Minimum Technical Standards (MTS)

|  |  |  |  |  | Perceived Artifacts |  |  |  | Visual Characteristics |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Statistic | Adequacy | Quality | Noise | Motion | Distortion | Aliasing | Blurriness | Contrast |
| Adherence | T2 | tau-b (SE) | - | 0.22(0.08)† | 0.19(0.07) | 0.1(0.07) | 0.17(0.06) | 0.14(0.06) | 0.23(0.09)† | 0.22(0.08)† |
|  |  | p-value | 0.227 | 0.010* | 0.012* | 0.133 | 0.005* | 0.017* | 0.008* | 0.009* |
|  | DWI | tau-b (SE) | - | -0.01(0.08) | -0.06(0.08) | -0.02(0.05) | 0.04(0.08) | 0.02(0.06) | -0.04(0.08) | -0.04(0.09) |
|  |  | p-value | 0.304 | 0.943 | 0.457 | 0.659 | 0.616 | 0.754 | 0.634 | 0.685 |

*Significant results with p-value ≤ 0.05

†Kendall Rank Tau-b ≥ 0.2 indicates presence of weak correlation

Table 4: Summary of results of statistical analysis of PIRADSv2 Minimum Technical Standards adherence association with image adequacy using Wilcoxon Rank Sum Test and with image quality and perceived artifacts using Kendall Rank Correlation.

| T2W Technical Standards | Adequacy† | Quality†† | Distortion†† |
| --- | --- | --- | --- |
|  | p-value | p-value | p-value |
| Slice Thickness | 0.436 | 0.276 | 0.310 |
| Slice Gap | 0.416 | 0.123 | 0.366 |
| FOV Upper Limit | 0.596 | 0.258 | *0.003 |
| IPRP | 0.970 | 0.885 | 0.528 |
| IPRF | 0.996 | 0.184 | 0.256 |
| Phase Encoding Dir. | 0.069 | *0.032 | *0.004 |
| DWI Technical Standards | p-value | p-value | p-value |
| High B value | 0.610 | 0.078 | 0.479 |
| TE | 0.123 | 0.628 | 0.866 |
| TR | 0.162 | 0.074 | *0.033 |
| Slice Thickness | 0.982 | 0.643 | 0.924 |
| Slice Gap | 0.930 | 0.347 | 0.473 |
| FOV Lower Limit | 0.777 | 0.741 | 0.982 |
| FOV Upper Limit | 0.864 | 0.354 | 0.446 |
| IPRP | 0.313 | 0.335 | 0.123 |
| IPRF | 0.313 | 0.335 | 0.123 |
| Phase Encoding Dir. | 0.576 | 0.350 | 0.779 |

*significant findings with p-value ≤ 0.5

†Chi-Square Test

††Wilcoxon Rank Sum Test

Abbreviations: FOV UL=field-of-view upper limit, FOV LL=field-of-view lower limit, IPRP=in-plane resolution phase, IPRF=in-plane resolution frequency

Table 5: Summary of association of adherence to individual PIRADSv2 Minimum Technical Standards and phase encoding direction with diagnostic adequacy, quality and distortion in prostate MRI as measured by the Chi-square and Wilcoxon Rank Sum tests for association.

|  |  |  | Perceived Artifact |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | tau-b (SE) | Noise | Motion | Distortion | Aliasing |
| Visual Characteristics | T2 | Quality | 0.58* (0.09) | 0.38 (0.07) | 0.31 (0.07) | 0.30 (0.07) |
|  |  | Blurriness | 0.58 (0.09) | 0.34 (0.07) | 0.25 (0.07) | 0.24 (0.06) |
|  |  | Contrast | 0.57 (0.09) | 0.30 (0.07) | 0.26 (0.07) | 0.25 (0.06) |
|  |  | % Reader Scores < 3 | 19% | 13% | 6% | 8% |
|  | DWI | Quality | 0.61* (0.09) | 0.31 (0.06) | 0.54* (0.09) | 0.40 (0.07) |
|  |  | Blurriness | 0.61 (0.09) | 0.29 (0.06) | 0.50 (0.08) | 0.34 (0.06) |
|  |  | Contrast | 0.60 (0.09) | 0.25 (0.05) | 0.46 (0.08) | 0.28 (0.06) |
|  |  | % Reader Scores < 3 | 30% | 10% | 27% | 26% |

Table 6. Kendall Rank Tau-b correlation statistics for Visual Characteristics with Perceived Artifacts and frequency of Reader Scores < 3 (interpretation difficult or impossible) for each type of artifact.