Characterization and correction of center-frequency effects in X-nuclear eddy current compensations on a clinical MR system

Mary A McLean (1,2), R Scott Hinks (3), Joshua D Kaggie (1), Ramona Woitek (1), Frank Riemer (4), Martin J Graves (1), Dominick J McIntyre (2), Ferdia A Gallagher (1), Rolf F Schulte (5)

(1) Department of Radiology, University of Cambridge, Cambridge UK

(2) Cancer Research UK Cambridge Institute, University of Cambridge, Cambridge UK

(3) GE Healthcare, Milwaukee, WI, USA

(4) MMIV, Department of Radiology, Haukeland University Hospital, Bergen, Norway

(5) GE Healthcare, Munich, Germany

*Correspondence to:

Dr Mary McLean

Cancer Research UK Cambridge Institute

Robinson Way

Cambridge CB2 0RE (UK)

# Email: Mary.McLean@cruk.cam.ac.uk

# Tel: +44 1223 769629

Running Title: X-nuclear eddy currents

Word count: 2704

Abstract

Purpose: The aim of the study was to investigate whether incorrectly compensated eddy currents are the source of persistent X-nuclear spectroscopy and imaging artifacts, as well as methods to correct this.

Methods: Pulse-acquire spectra were collected for 1H and X-nuclei (23Na or 31P) using the minimum repetition time (TR) permitted on a 3 T clinical MRI system. Data were collected in three orientations (axial, sagittal, coronal) with the spoiler gradient at the end of the TR applied along the slice direction for each. Modifications to system calibration files to tailor eddy current compensation for each X-nucleus were developed and applied, and data were compared with and without these corrections for: slice-selective MRS (for 23Na and 31P); 2D spiral trajectories (for 13C); and 3D cones trajectories (for 23Na).

Results: Line-shape distortions characteristic of eddy currents were demonstrated for X-nuclei, which were not seen for 1H. The severity of these correlated with the amplitude of the eddy current frequency compensation term applied by the system along the axis of the applied spoiler gradient. A proposed correction to eddy current compensation, taking account of the gyromagnetic ratio, was shown to dramatically reduce these distortions. The same correction was also shown to improve data quality of non-Cartesian imaging (2D spiral and 3D cones trajectories).

Conclusion: A simple adaptation of the default compensation for eddy currents was shown to eliminate a range of artifacts detected on X-nuclear spectroscopy and imaging.

Key words

magnetic resonance spectroscopy, MRI, carbon-13, phosphorus-31, sodium-23, X-nuclei, image artifacts, eddy currents

Introduction

Eddy currents arise when time-varying gradients induce currents in nearby conductive structures, such as coils and the cryostat, which oppose the change in gradient field, and these currents are frequently responsible for spectroscopic and imaging artifacts. These spatially-varying eddy currents accumulate due to time-varying gradients, which can be characterized by multiple components of exponential time-constant terms. One component of an eddy current at observation time t can be expressed as:

et= -0tα⋅dg/dt'⋅exp-βt-t'dt' 			[1]

where dg/dt’ is the rate of change of a gradient at time t’, α is an amplitude term, 1/β is the time constant of the eddy current, and (t-t’) is the delay between eddy current generation and observation. When the gradients are at zero or constant amplitude, the eddy currents consist of multiple components that each decay exponentially with different time constants:

et=etR⋅exp⁡(-βt-tR)		[2]

where tR is the end of the gradient ramp.

These eddy currents generate magnetic fields, which can most easily be observed by recording an FID after the gradient has been switched off. The eddy currents cause a phase accumulation (Δφ) between adjacent points, from which the eddy current field B can be calculated:

Δφ/Δt=ω= γB					[3]

The unwanted effects of eddy currents on MR experiments can be counteracted by system modification to prevent or counteract their production, by post-processing to remove the effects, or by a combination of these methods.

A highly effective post-processing correction for eddy currents is available for 1H-MRS using a separately acquired reference spectrum from the unsuppressed water to deconvolve the phase of the signal (1). This technique is less applicable in X-nuclear spectroscopy, where there is no equivalent of the unsuppressed water peak as a basis for correction. Collection of separate data solely for eddy current correction would be possible (2), but would either add to the scan time significantly for low-SNR X-nuclei, or would involve extrapolating corrections for X-nuclei from 1H data, which may be inexact. For imaging experiments, the actual k-space trajectories played out can be measured and compensated (3), but that must be repeated for each different type of imaging experiment.  Finally, collection of dynamic calibration data for post-processing correction using field cameras can be employed (4), but this requires additional hardware.

Most modern MR scanners include extensive methods to reduce the effects of eddy currents during acquisition. On some scanners, principally preclinical devices, this involves the use of dynamic Bn coils, which produce an eddy current compensation that is insensitive to the gyromagnetic ratio (5). On others, such as General Electric (GE) devices, the gradient and frequency terms are compensated separately as follows. A calibration procedure with a complex arrangement of small phantom samples and detector coils is carried out (a field camera as mentioned above), where eddy currents are measured directly by acquiring data at a range of delays during and after application of a gradient (6).  The measurements are made with gradients applied along the 3 orthogonal axes (X, Y and Z), with multiple coils co-located with water samples at locations offset from the isocenter along these directions, and the eddy currents are decomposed into spatially dependent and independent effects. For example, if six coils are used to acquire data Bn(t) where n = 1:6, these components can be determined by a least squares fit to the data:

B1tB2tB3tB4tB5tB6(t)=B0t+Gx(t)⋅X1X2X3X4X5X6+ Gy(t)⋅Y1Y2Y3Y4Y5Y6+ Gz(t)⋅Z1Z2Z3Z4Z5Z6		  [4]

Following this decomposition, the B0 and linear effects are directly fitted with multiple exponential functions, each with a time constant and amplitude term. The spatially dependent effects caused by the eddy currents are compensated with an applied equal and opposite pre-emphasis to the gradient waveforms along X, Y and Z. Similarly, the spatially independent but time-varying field ΔB0(t) is compensated by modulating the centre frequency:

∆f0t=γ2π∆B0t 					[5]

The experiments are repeated with these corrections applied, new fits are performed to the residual eddy currents measured, the correction terms are modified, and this iterative cycle is repeated until the residual eddy currents are below an acceptable threshold.

Although the eddy currents themselves are independent of field strength, precession frequency is proportional to the gyromagnetic ratio, so it is important that any frequency correction should take this into account. In our experience, visual inspection of data has often suggested a possible contribution from eddy currents to the poor quality of X-nuclear spectra and images. Spectral peaks have been distorted, and images have been distorted or displaced relative to anatomic 1H scout images. However, subsequent 1H-MRS failed to demonstrate a problem, despite its sensitivity to eddy current effects in general (1).  It is therefore likely that eddy current effects on X-nuclei may be overlooked on human imaging systems where the focus is optimizing 1H-MRI quality for routine clinical use, and that some modifications for different gyromagnetic ratio may be needed for optimal acquisition.

In this paper, we present a method to characterize the severity of eddy current artifacts on different nuclei, and we demonstrate a fix on a clinical system which greatly improves the quality of X-nuclear images and spectra.

Methods

Experiments were performed on a 3 T MR750 MR system (software version DV26.0_R02, GE Healthcare, Waukesha WI).  Acquisition was performed with quadrature birdcage head coils (Rapid Biomedical, Rimpar, Germany) dually resonant for 1H (128 MHz) and either 23Na (33.9 MHz), 13C (32.1 MHz) or 31P (51.7 MHz).

A 16-cm diameter MRS head sphere phantom (GE Healthcare, Waukesha WI) was used for 1H, 23Na and 31P spectroscopy. For 13C, a 3-cm diameter sphere of 1 M 13C-labelled bicarbonate was used (GE Healthcare, Waukesha WI), because natural abundance phantoms would have given low signal, often complicated by J-coupling with nearby 1H nuclei. To assess the performance of sodium 3D cones acquisition, a gradient calibration phantom (DQA III, GE Healthcare, Waukesha WI) was emptied and refilled with 80 mM sodium chloride solution.

Slice-selective pulse-acquire spectra with matched parameters were collected for 1H and X-nuclei (23Na, 31P, and/or 13C) using the minimum allowed TR (420 ms) and with incrementally increased delays between the large spoiler gradient at the end of one TR period and the beginning of the next. Data were collected in three orientations (axial, sagittal, coronal) with the spoiler gradient at the end of the TR applied along the slice direction. The spoiler gradient amplitude was 22 mT/m, pulse width 6.5 ms, ramp 0.42 ms, area 0.152 mT s m-1.

As described in the Introduction, eddy current correction on this system involves applying pre-emphasis to the frequency and the gradient waveforms using multiple exponential terms, with a time constant and amplitude for each. The amplitude terms (α) of the frequency modulation for eddy current correction were altered in the following ways by creation of modified system calibration files: they were set to zero, or they were rescaled for the gyromagnetic ratio of each X-nucleus according to:

αX = α1H * γ X / γ1H 				[6]

Swapping between these system calibration files was performed within an examination by the execution of a Python script. Data were collected and compared with these different calibration files using slice-selective MRS, 2D spirals (7) (single-shot 60-point spiral, FOV 24 cm, nominal pixel size 4 mm, 2 cm thick slice, 90° flip angle, 128 averages, TR = 75 ms, 0:10 min), and 3D cones (8,9) (24 cm FOV, nominal pixel size 3 mm isotropic, TR = 45.6 ms, TE = 0.46 ms, hard pulse flip angle 90°, 1 average, 2552 interleaves, 1:56 min).

Spectra and images were reconstructed and displayed in MATLAB (The Mathworks, Natick MA) and SAGE (GE Healthcare, Waukesha WI). They were assessed qualitatively for the presence and severity of eddy current artifacts, and the size of the eddy current induced frequency shift was measured as the displacement in Hz between the peak in the last transient in a series of 8 spectra, which was distorted by the spoiler gradients in the previous repetitions, and the first spectrum, which was not preceded by a spoiler gradient and was therefore unaffected.

Results

In pulse-acquire spectroscopy at the minimum TR, small distortions due to eddy currents were seen in 1H MR spectra of water acquired in axial and coronal orientations (with crusher gradients applied along Z and Y, respectively), but not when the crusher gradients were applied along X (sagittal orientation; Fig 1). When spectra were acquired from the same phantom at 23Na frequency, there was still no evidence of eddy currents in the sagittal orientation, but the distortion was greater than seen for 1H in the coronal orientation, and very much greater than 1H in the axial orientation.

To investigate this further, we modified the system calibration files such that the gradient pre-emphasis was still applied but the frequency modulation amplitude terms were set to zero. Pulse-acquire spectroscopy then showed larger eddy current effects on 1H than on X-nuclei, particularly in the Z direction (Fig 2).

An example of how these relative shifts varied with increasing delay between the spoiler pulse at the end of one TR period and the start of the next is shown in Figure 2B-C.  The rate of decay with increasing time delay appeared very similar although not identical between 1H and 23Na, and the magnitude of the frequency shifts was roughly in proportion to the gyromagnetic ratios. Although the frequency shifts with zeroed correction were largest in Hz for 1H, when rescaled to ppm they were similar between nuclei. Applying default frequency corrections for X-nuclei resulted in clear overcompensation, as the magnitude of frequency shift became much larger and changed sign (Fig 2C).

We further confirmed that the existing frequency modulation produced an overshoot for a different X-nucleus (phosphorus), and that a simple scaling of the amplitude coefficients for modulation could dramatically improve the data (Fig 2D and E). In Fig 2D, the existing correction worked well for 1H: with zeroed frequency modulation, severe distortions were seen and the tallest peak in the magnitude spectrum was shifted by +7.0 Hz; when the default 1H-optimized frequency modulation was applied, the lineshape was much improved and the shift was 0.0 Hz. For the phosphorus acquisition (Fig 2D), the zeroed modulation produced a shift in the same direction as 1H but smaller, +3.1 Hz; the default 1H-optimized f0 correction produced a shift in the opposite direction, -3.1 Hz; and the f0 correction with amplitudes rescaled for 31P gave a shift of 0.0 Hz and an undistorted lineshape.

The original impetus for exploring the eddy current behaviour had been the observation that low frequency, non-Cartesian X-nuclear images were consistently displaced superiorly relative to an anatomic 1H image.   This is shown in Figure 3 with coronal spiral images, using a spherical phantom of labelled 13C-bicarbonate. When the rescaled f0 correction terms are used for compensation, the displacement is reduced.

Similar displacement effects were observed when imaging sodium using 3D non-Cartesian sequences such as cones. Additionally, when using a short TR there was considerable blurring and distortion: rescaling the eddy current f0 correction terms for sodium was shown to improve image quality markedly in a 3D cones dataset (Fig 4).

Discussion

Spectroscopy using a spoiler gradient and the minimum allowed TR presents a simple, quick and effective method to check the relative performance of eddy current compensation for different nuclei and axes. This approach is less sensitive and precise than the methods used to calibrate the system for 1H: those experiments employ very long gradient pulses to diminish the interference between eddy currents caused by turning the gradients off and those caused by turning them on. Common 1H methods also use a field camera, whereas for many X-nuclei the SNR obtainable in a rig of tiny samples might be insufficient for full multi-exponential fitting. The advantage of the method employed in the current study is that it can be performed without specialist equipment and without changing the default pulse sequence variables, which should increase the robustness for inter-site comparisons.  As shown in Figure 1, the method still retains sufficient sensitivity to detect and characterize problems.

We demonstrated that on a GE MR750 MRI scanner, eddy current compensation was significantly worse for X-nuclei than for 1H, which is likely to be due to the applied frequency modulation for eddy currents. Since this modulation was the same for all nuclei regardless of their resonant frequency, it is expected to over-correct the lower-frequency X-nuclei. We have confirmed that zeroing this frequency modulation produced lineshape distortions and frequency shifts for 1H which were greater than for the X-nuclei (Fig 2). Application of the default 1H-optimized f0 correction compensated the 1H acquisition adequately, but produced an overshoot in the frequency shift of 31P and 23Na. Correction of this error was demonstrated using a simple rescaling of a series of amplitude coefficients in an eddy current compensation file by the ratio of gyromagnetic constants for the X-nucleus relative to hydrogen. A separate correction must be used for each nucleus since the degree of over-compensation is  dependent on the gyromagnetic ratio. We predict that this effect would be minor for 19F because of its similar frequency to 1H, and that artifacts would worsen progressively for 31P, 23Na, 13C, and 2H. It should be noted that this error is specific to the method of f0 eddy current compensation used: it would not be expected to occur where the correction was applied through the use of zeroth order field coils, which we believe to be more common on pre-clinical systems.

Although it appeared that the time constants as well as amplitude coefficients of the eddy current terms might vary slightly between nuclei (Fig 2B), these apparent differences were probably caused by the imprecision of quantifying small frequency shifts by fitting a single frequency to broad peaks using limited spectral resolution. This limitation prevents confident multi-exponential fitting of the curves produced, which could in theory be used to re-derive the f0 correction parameters for each X-nucleus individually. Additionally, the artifacts we observed were empirically well corrected by adjusting amplitude coefficients only (Fig. 2C and 2E); therefore, the time constants were assumed to be the same for all nuclei.

Application of these corrections was shown to improve the quality not only for spectroscopy, but also for fast imaging of X-nuclei using spiral and 3D cones sequences. These techniques are widely used in clinical research, particularly for 23Na (10) and hyperpolarized 13C (11) and 129Xe (12). Overcompensation of eddy current f0 terms may have had quite widespread effects on signal localisation and quantification using these approaches. Other fast imaging sequences such as EPI would also be expected to suffer from overcompensated eddy currents, but they are not implemented for X nuclei on our system and were not tested. Diffusion-weighted (DW) MRS of X-nuclei on clinical systems (13) would be expected to be especially sensitive to eddy currents: a pulse sequence to perform DW-MRS was implemented in 2014 on our system but abandoned due to the problems experienced in particular when applying gradients along the Z axis for X-nuclei. This is consistent with the effects seen in Figure 1 and the relative eddy current correction coefficients on our system.

The results reported here are limited to a single MR scanner from a single manufacturer (GE), and it seems that not all GE scanners are equally affected. Lu et al (3) concluded that since 1H and 23Na were affected approximately equally by eddy currents on their GE 3T scanner, they could correct 23Na based on 1H calibrations. However, there are indications that some scanners from other manufacturers may also be affected (14).

It should be noted that the experiments which show this artifact most clearly are not often performed as part of human studies. MR spectroscopy is rarely acquired at the minimum possible TR, because T1s of metabolites of interest tend to be long, such that SNR per unit time is maximized by acquiring at longer TRs. Inversion-recovery MRS experiments with short inversion times are more commonly undertaken, given their usefulness in estimating T1, but few clinical studies of quantitative T1 relaxometry have been reported on X-nuclei. Additionally, the error is most obvious when the uncorrected Δf0(t) term is large to begin with (Fig 1). Where the frequency shifts due to eddy currents are small on 1H, even a four-fold overcompensation (as is the case for 13C, where the gyromagnetic ratio is four times smaller) may produce only subtle effects.

A temporary solution of tailoring calibration files for each nucleus was demonstrated to completely eliminate evidence of eddy current artifacts on a clinical system; however, this carries a risk to the clinical service if operators fail to reset these for subsequent 1H acquisition. Automation of this fix within the system applications software would be desirable, to eliminate the need for user intervention. We have demonstrated that such a fix has the potential to markedly improve multi-nuclear image quality.

Acknowledgements

This study was funded by Cancer Research UK (RG86786/A25117), National Institute of Health Research-Cambridge Biomedical Research Centre, and Addenbrooke’s Charitable Trust.

References

1. Klose U. In vivo proton spectroscopy in presence of eddy currents. Magn Reson Med 1990; 14:26-30.

2. Boesch CH, Gruetter R, Martin E. Temporal and spatial analysis of fields generated by eddy currents in superconducting magnets: optimization of corrections and quantitative characterization of magnet/gradient systems. Magn Reson Med 1991; 20: 268-284.

3. Lu, A, Atkinson IC, Claiborne TC, Damen FC, Thulborn KR. Quantitative sodium imaging with a flexible twisted projection pulse sequence. Magn Reson Med 2010; 63: 1583-1593.

4. Barmet C, De Zanche N, Pruessmann KP. Spatiotemporal magnetic field monitoring for MR. Magn Reson Med 2008; 60: 187-197.

5. Fry ME, Pittard S, Summes IR, Vennart W, Goldie FTD. A programmable eddy-current compensation system for MRI and localized spectroscopy. JMRI 1997; 7: 455-458.

6. Glover GH, Pelc NJ. Method for magnetic field gradient eddy current compensation. US Patent Number 4698391 (1987).

7. Wiesinger F, Weidl E, Menzel MI, Janich MA, Khegai O, Glaser SJ, Haase A, Schwaiger M, Schulte RF. IDEAL spiral CSI for dynamic metabolic MR imaging of hyperpolarized [1-(13)C]pyruvate. Magn Reson Med 2012; 68:8-16.

8. Gurney PT, Hargreaves BA, Nishimura DG. Design and analysis of a practical 3D cones trajectory. Magn Reson Med 2006; 55:575-582.

9. Riemer F, Solanky BS, Stehning C, Clemence M, Wheeler-Kingshott CA, Golay X. Sodium (23Na) ultra-short echo time imaging in the human brain using a 3D-Cones trajectory. MAGMA 2014; 27:35-46.

10. Thulborn KR. Quantitative sodium MR imaging: A review of its evolving role in medicine. Neuroimage. 2018;168:250-68.

11. Kurhanewicz J, Vigneron DB, Ardenkjaer-Larsen JH, Bankson JA, Brindle K, Cunningham CH, Gallagher FA, Keshari KR, Kjaer A, Laustsen C, Mankoff DA, Nelson SJ, Pauly JM, Lee P, Ronen S, Tyler DJ, Rajan SS, Speilman DM, Wald L, Zhang X, Malloy CR, Rizi R. Hyperpolarized 13C MRI: path to clinical translation in oncology. Neoplasia 2019; 21(1): 1-16.

12. Ebner L, Kammerman J, Driehuys B, Schielber ML, Cadman RV, Fain SB. The role of hyperpolarized 129xenon in MR imaging of pulmonary function. Eur J Radiol 2017; 86: 343-352.

13. Koelsch BL, Reed GD, Keshari KR, Chaumeil MM, Bok R, Ronen SM, Vigneron DB, Kurhanewicz J, Larson PEZ. Rapid in vivo apparent diffusion coefficient mapping of hyperpolarized 13C metabolites. Magn Reson Med 2015; 74:622-633.

14. McLean MA,  Hinks RS, Kaggie J, Woitek R, Riemer F, Graves MJ, McIntyre DJO, Gallagher FA, Schulte RF. Characterization and elimination of X-nuclear eddy current artifacts on clinical MR systems. Proc ISMRM 2020; 3385.

Figure Legends

Figure 1: Eddy current distortions in a phantom spectrum acquired using 1H-MRS and 23Na-MRS.  Slice-localized MRS was acquired along each orthogonal axis in a large spherical saline phantom (5000 Hz, 2048 points) with application of the default correction for eddy currents. While only minor distortion is seen in the axial orientations on 1H (left), marked distortions typical of eddy currents are seen in a 23Na spectrum from the same slice (right), while there is less distortion for 23Na in the coronal orientation, and virtually none in the sagittal orientation.

Figure 2: Eddy current distortions with frequency modulation disabled and modified. (A) A stack of 1H-MR spectra acquired with zeroed frequency modulation, in an axial orientation at the minimum TR (420 ms). The first acquisition in the series is free of distortion, while subsequent acquisitions show frequency shifts and phase distortion caused by the spoiler gradients applied in the previous TR periods. (B) The frequency shift (ppm) between the tallest peak in the first and last spectra in a stack as shown in (A) was measured at a range of time delays between the end of the spoiler gradient in one TR period and the excitation in the next, for both hydrogen and sodium with zeroed f0 correction terms, and for hydrogen with the default correction applied. (C) As in (B), frequency shifts are shown for sodium with the default f0 terms and with the corrected terms recalculated for the gyromagnetic ratio. (D) 1H-MRS acquired using a 1H/31P head coil, with either zeroed frequency correction (dotted line) or default 1H-optimized correction (solid). (E) 31P-MRS acquired with the same coil and phantom, using frequency corrections: zeroed (dotted line); default 1H-optimized (thick line); custom 31P-optimized (narrow line).

Figure 3: Spatial mis-registration of spiral X-nuclear images. With the default eddy current correction settings (left), a coronal spiral image of a spherical phantom of 1 M 13C-bicarbonate (overlaid in color) is seen to be offset in the superior-inferior direction relative to the corresponding 1H-MRI image. With corrected eddy current f0 terms (right), this offset is reduced. For clarity, only the central 6 cm of a 24-cm FOV are shown.

Figure 4: Distortion of X-nuclear cones images. (A) With the default eddy current correction settings, a series of 23Na axial 3D cones images of a phantom of 80 mM saline acquired at short TR (46 ms) is seen to be blurred and distorted. This is particularly evident in the coronal (B) and sagittal (C) projections. With corrected eddy current f0 terms (D,E,F), the blurring and distortion are greatly reduced.