Erythrocyte transketolase activity coefficient (ETKAC) assay protocol for the assessment of thiamine (vitamin B1) status

Kerry S. Jonesa, Damon A. Parkingtona, Lorna J. Coxa,b, Albert Koulmana

a NIHR BRC Nutritional Biomarker Laboratory, MRC Epidemiology Unit, University of Cambridge, Cambridge, UK

b MRC Elsie Widdowson Laboratory, Fulbourn Road, Cambridge, UK*

* (closed down in December 2018)

Corresponding author contact information:

Kerry S. Jones

NIHR BRC Nutritional Biomarker Laboratory,

MRC Epidemiology Unit,

Clifford Allbutt Building,

Cambridge Biomedical Campus,

Hills Road,

Cambridge,

CB2 0AH

UK

Email: kerry.jones@mrc-epid.cam.ac.uk

Short title: ETKAC assay protocol

Key words: thiamine; vitamin B1; ETKAC; beriberi; TPPE

Abstract

Vitamin B1 (thiamine) is an essential nutrient that acts as a cofactor for a number of metabolic processes, particularly in energy metabolism. Symptoms of classic thiamine deficiency are recognised as beriberi, although clinical symptoms are non-specific and recognition of sub-clinical deficiency is difficult. Therefore, reliable biomarkers of thiamine status are required.

Thiamine diphosphate (ThDP) is a cofactor for transketolase including erythrocyte transketolase (ETK).The ETK activity assay as an indirect, functional marker of thiamine status, has been used for over 50 years. The ETK activity assay provides a sensitive and specific biomarker of thiamine status, however, there is a lack of consensus over the cut-offs for deficiency, partly due to a lack of assay harmonization.

Here, we provide a step-by-step protocol for the measurement of ETK activity and calculation of the ETK activity coefficient (ETKAC), including detailed explanations of equipment and chemicals required and guidance for quality control procedures.

Harmonization of protocol will provide the basis for the development of internationally recognized cut-offs for thiamine insufficiency. The establishment of quality control materials and a quality assurance scheme are recommended to provide reliability. This will ensure the ETK activity assay remains an important method for the assessment of thiamine status.

- Introduction

Thiamine (vitamin B1) is an essential micronutrient. In humans, the majority of thiamine exists in phosphorylated forms: thiamine monophosphate (ThMP), thiamine diphosphate (ThDP) (also known as thiamine pyrophosphate (ThPP)) and thiamine triphosphate (ThTP). ThDP is the major form, making up around 80% of total blood thiamine and is mainly present in erythrocytes and white blood cells.

ThDP is an essential cofactor for a number of metabolic enzymes, particularly those associated with oxidative and non-oxidative carbohydrate metabolism. These include pyruvate dehydrogenase, responsible for the conversion of pyruvate to acetyl coenzyme A, α-ketoglutarate dehydrogenase in the Krebs cycle and the branched chain α-keto acid dehydrogenase complex.  In non-oxidative carbohydrate metabolism, ThDP is a cofactor for the transketolase in the pentose phosphate pathway, necessary for the production of pentoses and NADPH used in nucleic acid and fatty acid synthesis, respectively.1

Thiamine deficiency is associated with non-specific clinical symptoms affecting the cardiovascular, muscular and nervous  systems. The classical manifestation of thiamine deficiency is beriberi. The populations most at risk are breastfed infants of thiamine-deficient mothers in low-income countries, particularly those with poor diet diversity and where rice is the staple food.2,3 In high-income countries, deficiency is rare but may be present in the elderly, chronic alcoholics and persons with acute or chronic medical conditions associated with malnutrition.1 In addition to the effect of acute deficiency, evidence also exists that sub-clinical deficiency may have long-term effects in relation to cognition and gross motor skills.2

In order to investigate the potential impact of sub-clinical deficiency and to improve our understanding of the global prevalence of thiamine deficiency, biomarkers to assess thiamine status are essential. The measurement of serum or plasma thiamine concentration is of limited use since it represents only a small portion of thiamine in the blood and is affected by a number of disease states.1 Although a correlation between plasma thiamine and erythrocyte ThDP has been shown4, these data were from a thiamine fortification trial; plasma thiamine is affected by recent dietary intake and is therefore not suitable as a biomarker of long-term thiamine status. Urinary thiamine excretion may correlate with intake at all but low thiamine intakes. However, it does not provide information on thiamine status particularly when status is low as thiamine may be conserved. In addition, single 24-hour urine collections are of limited use due to large within-subject variability. 5 Longer-term thiamine status from human blood samples can be assessed either by direct measurement of ThDP in erythrocytes or whole blood, or by the measurement of the degree of ThDP-saturation of erythrocyte transketolase (ETK), a ThDP-dependent enzyme. The advantages and disadvantages of both the direct and indirect assays were reviewed recently 1,2,6, including a detailed description of the direct measurement of ThDP. However, the methodology to measure the ETK activity has received less attention.

ETK activity is commonly expressed as a ratio or ‘activity coefficient’ (ETKAC). Two measurements are made to determine the EKTAC: (1) the basal activity of ETK and (2) ‘activated’ or ‘stimulated’ ETK activity by addition of exogenous, excess ThDP. The ratio of the activated to basal activity gives the ETKAC and provides a proxy measure for the in vivo activity of ETK and an indication of the availability of ThDP and thiamine status.

In this manuscript, we describe the background to the ETK activation assay and discuss aspects relating to reliably and robustly performing, reporting and interpreting the ETK activity data. In addition, we present a step-by-step protocol that provides detailed guidance and should enable laboratories to set up the method independently and improve harmonization.

- Historical background

The ETK activity assay in its early form for the application of thiamine deficiency assessment was established by Brin 7 and was later improved and adapted as understanding and technology progressed. Assays moved from the less-specific colorimetric determination of substrate disappearance or appearance of end-products to UV detection of NADH. 8 The assay relies on the detection of small changes in enzyme activity and early methods lacked the necessary sensitivity and had relatively poor precision. The assay was further refined and adapted for semi-automated clinical chemistry analysers. This improved measurement precision, even at observed high absorbance, and the ability to run basal and activated measurements side-by-side further improved performance.9,10 For additional convenience and to remove the need for large, expensive equipment, the assay has been further adapted to a 96-well plate format and measured with a plate reader11. The assay as it is currently performed in our laboratory is described herein.

- Chemistry of assay

The ThDP-dependent enzyme ETK catalyses the metabolism of pentoses in two reactions in the pentose phosphate pathway:

- xylulose-5-phosphate + ribose-5-phosphate  sedoheptulose-7-phosphate + glyceraldehyde-3-phosphate

- xylulose-5-phosphate + erthyose-4-phosphate  fructose-6-phosphate + glyceraldehyde-3-phosphate

In the current assay, the resultant glyceraldehyde-3-phosphate is metabolised to dihydroxyacetone phosphate, which in the final reaction is reduced to glycerol-3-phosphate by glycerol-3-phosphate dehydrogenase and NADH (Figure 1). The assay monitors the rate of oxidation of NADH by measuring the decrease in absorbance at 340 nm. In the described protocol, xylulose-5-phosphate is not required as a substrate; ribose-5-phosphate is converted to xylulose-5-phosphate by erythrocyte thiamine-independent enzymes ribosephosphate isomerase and ribulosephosphate-3-epimerase. 8

- Calculation and interpretation of ETKAC

Calculation of ETKAC involves in erythrocyte lysates: (1) measurement of the basal activity of ETK; (2) measurement of ‘activated’ or ‘stimulated’ ETK activity by addition of exogenous, excess ThDP. The basal test represents the active holo-transketolase activity. Addition of exogenous ThDP activates apo-transketolase and the ‘stimulated’ activity represents the activity of both the (now-activated) apo-transketolase and holo-transketolase. Takeuchi et al., demonstrated in experiments with apo-transketolase enzymes, that ETKAC reflects the degree of saturation of ETK with ThDP. 12

The results of the basal and stimulated activities are commonly presented as a ratio or activity coefficient, the ETKAC calculated as stimulated activity/basal activity. Alternatively, results may be presented as a percentage [% activation  = (ETKAC x 100) – 100] (also known as thiamine pyrophosphate effect (TPPE)).

In thiamine sufficiency, addition of exogenous ThDP will make little difference to the enzyme activity and hence an ETKAC close to 1 will be obtained.  In thiamine insufficiency or deficiency addition of exogenous ThDP has a progressively greater effect on ETKAC to provide a continuum of thiamine status.13

Although there is no international consensus on cut-offs, the commonly used threshold for risk of deficiency is an ETKAC of >1.25.13,14 An ETKAC of less than 1.15 indicates sufficiency and values between 1.15 and 1.25 suggest a low risk of clinical deficiency. 13 Others have suggested values of ≥1.2 indicate deficiency.15 Beriberi is typically associated with ETKAC values >1.4.2

The use of  ETKAC rather than absolute measures of ETK activity is preferred for three main reasons:  (1) the between-subject variation in basal activity is large; (2) it is assumed, but not certain, that apo-enzyme levels are not affected by vitamin deficiencies 5 and (3) it reduces the need for the precise definition of assay conditions (e.g. optical path length), which are critical to calculate absolute enzyme activities.16 Others have suggested using combinations of the stimulated ETK activity and the ETKAC in interpretation.17

However, the use of basal ETK activity rather than ETKAC has also been recommended, particularly for the assessment of thiamine status in infants due to potential low levels of apo-enzyme caused by long-term, low exposure to thiamine. 18 A recent study of children and adults found that whilst both basal and ETKAC correlated with erythrocyte ThDP the correlation was stronger for basal activity.19 Basal (or activated ETK) activity needs to be expressed per mass unit of hemoglobin and it is necessary to measure the hemoglobin concentration in the same lysate as used for the ETK activity assay. In addition, it is necessary to calculate the factor needed to convert rate to activity.

- Recommended sample type

Washed erythrocytes are the recommended sample type. Leucocytes have relatively large amounts of thiamine and transketolase activity can therefore be considerably higher in whole blood compared to erythrocytes.8,20 The use of leucocytes 21,22 to assess transketolase activity has been reported but has not been applied routinely, possibly because the shorter half-life of leucocytes compared to erythrocytes would result in a only short-term biomarker of thiamine status.

The use of whole blood, a more convenient sample compared to washed erythrocytes, has also been considered. Whilst transketolase activity was higher in whole blood, the transketolase activity coefficient was similar between whole blood and erythrocyte hemolysates.20,23 However, there are concerns around the use of whole blood, partly due to between-person variability in leucocyte count, particularly in clinical conditions that may impact leucocyte number. 24

Whole blood collected in blood tubes containing either lithium heparin (LH) or ethylenediaminetetraacetic acid (EDTA) anti-coagulant can be used as the starting sample for the preparation of washed erythrocytes. Agreement in our own laboratory between the two sample types is shown in Figure 2.

- Sample stability

A small number of studies have investigated the stability of ETK activity and ETKAC across a number of analytical stages including in whole blood and hemolysates, and at a range of temperatures from refrigeration at 4 °C to frozen storage at -20 °C and -70 °C, possibly reflecting the availability of cold storage in different clinical and field site settings.

Whole blood

In whole blood kept at 4 °C or at room temperature, ETKAC was stable for up 24 h, decreasing thereafter; 25,26 however, basal activity was reportedly stable for up to 4 d.25

In our laboratory, whole blood samples from 15 participants kept at 4 °C were also observed to be stable for 24 h with no significant change in ETKAC observed in LH whole blood (% geometric mean change (95% confidence interval) -0.3 (-2, 2)) or  EDTA whole blood (-1 (-3, 1) and showed good agreement (Figure 3).

Washed erythrocytes

Puxty et al., observed greater deterioration of ETKAC after storage of washed erythrocytes at -20 °C in participants in an ‘insufficient’ group (ETKAC >1.15) group than in a sufficient group (ETKAC <1.15) and speculate this may be due to the stability of different isoenzymes. In the sufficient group, samples were stable for up to 14 days at -20 °C but for only 4 d in the ‘insufficient’ group.26 Washed erythrocytes stored at 4 °C were stable for less than 2 d.26 Other studies observed no consistent change over time in basal and stimulated activity.9

Hemolysates

Hemolysates can be kept frozen at -20 °C for at least 4 weeks, however caution should be paid to freeze-thaw cycles since reports suggest significant falls in enzyme activity can occur after two freeze-thaw cycles.8 Another study showed a decrease of 7% in ETKAC in hemolysates analysed after freezing at -18 °C for 2 weeks compared with fresh samples.27 As with washed erythrocytes, samples kept at -70 °C remained stable for at least 2 months.25 In our own laboratory, we observed a less than 1% change in ETKAC in a set of 59 hemolysates stored at -70 °C for 6 months and excellent agreement across a range of EKTAC (Figure 4). Enzyme activities reportedly remain stable at -70 °C for over one year.2 Storage of washed erythrocytes and hemolysates at -70 °C or below is recommended.16

- Assay and individual variability

Analytical variability

Our routine method includes 3 internal QC materials with ETKAC (mean (SD)) of 1.07 (0.03), 1.15 (0.05) and 1.18 (0.05) and corresponding between batch %CVs of 3.1, 3.9 and 4.1% (n 25). Targets of <5% are recommended.

Within-individual variation

A study from 1990 reported within-individual variation of basal and activated transketolase activity between 3% and 6% measured in 20 samples collected from each of 4 individuals over 60 days on a constant diet.28

Between-individual

ETKAC data from the UK National Diet and Nutrition Survey (NDNS) Rolling Programme (Years 7  - 8 (2014 – 2016)29 are presented in Table 1 and show the range of values in a largely thiamine-replete population. Between-individual CV is around 5%. The 2.5th – 97.5th percentile range in the UK population based on the NDNS Rolling Programme data for Years 1 – 8, using the ETKAC method described herein, is 1.01 to 1.21.

- Limitations and other factors

Genetic differences leading to different transketolase isoforms that may have different activity or stability,30 or be associated with the manifestation of thiamine deficiency in Wenicke-Korsakoff syndrome, have been suggested.30,31 However, other studies have found no evidence for isoforms or disease associations and suggest earlier observations were due to the methodologies used.32–36

A venous whole blood sample is required for the determination of ETKAC (and whole blood ThDP). Whilst a venous blood sample is a common sample type, collection and processing of blood samples in field- or limited resource-settings presents a number of challenges including laboratory infrastructure for washing of erythrocytes and cold storage. Capillary blood sampling and dried blood spots (DBS) provide a more convenient sample. Recently, Huang et al. described a DBS method for ThDP which showed a high degree of correlation with whole blood results (r 0.964; p < 0.0001) although ThDP on DBS was not stable at room temperature.37 Further work is required to determine the feasibility of measuring ETKAC in whole blood as either venous sample or DBS.

The impact of clinically observed thiamine deficiency, ThDP concentration, their effect on transketolase enzyme and consequences for interpretation of transketolase activity have been a topic of discussion for many years.17,38 It has been suggested that liver disease may directly affect apo-transketolase concentration causing apparently normal ETKAC in the presence of clinical symptoms of thiamine deficiency.  However, other studies posit that the decrease in apo-transketolase is a response to thiamine deficiency. 17,38 Either way, chronic thiamine deficiency may cause loss of apo-transketolase and evidence suggests this is due to an effect on synthesis rather than catabolism of transketolase. 39  The impact of this effect on interpretation of the EKTAC is however uncertain. Price et al., observed a lower activation after in vitro stimulation of ETK compared with in vivo activity following thiamine supplementation that may be related to irreversible inactivation of the enzyme in vitro. 40 Thus, observations in vitro may not represent the in vivo situation.  Clinical conditions, e.g. certain anaemias or diseases that affect erythrocyte cell survival time and the age distribution, may also affect interpretation of ETKAC. Holo-transketolase enzyme concentration is lower in older than younger erythrocytes 41 and as holo-transketolase decreases there is a consequent increase in the ETKAC.26,41

- Conclusions

The ETKAC assay provides an indirect, functional measure of thiamine status. Perceived limitations of the assay include relatively poor precision, lack of standardization, instability of the transketolase enzyme and lack of consensus over cut-offs for deficiency.42 However, these limitations equally apply to other assays of thiamine status, such as whole blood ThDP. ETKAC has advantages over methods for the direct measurement of ThDP, including the relative ease and requirement for less specialized equipment making the assay potentially more affordable and sustainable. Furthermore, ETKAC provides a potentially better marker of long-term thiamine status that is less affected than ThDP by acute changes in thiamine intake.43

In this manuscript, we provide a detailed protocol for the measurement of ETK activity that can be used as a foundation for assay harmonization. There is also a need to develop an external quality assurance scheme to provide for the independent monitoring and assessment of laboratory performance. The establishment of a network of international laboratories willing to share samples for an inter-laboratory comparison could provide a first step in the development of an EQA scheme. The difficulty of obtaining blood samples in sufficient volume from thiamine-deficient individuals results in a lack of quality control materials that include the thiamine-deficient range. Therefore, there is a need to develop quality control and reference materials with consensus values spanning both sufficiency and deficiency, which can be supplied to laboratories in order to monitor, understand and standardise assay performance.

Improvements in harmonization will also provide opportunity for investigating outstanding questions related to thiamine physiology and EKTAC assays, including the establishment of cut-offs for deficiency, characterising agreement with other markers of thiamine status (e.g. ThDP) and to better understand conditions where interpretation of ETKAC may be confounded.

Acknowledgements

KSJ and DAP wrote the manuscript; LJC designed the assay whilst at MRC EWL (UK Medical Research Council programme U105960384); LJC and DAP performed the experiments; KSJ, DAP, LJC and AK reviewed and approved the manuscript and accept responsibility for the contents.

KSJ, DAP and AK are supported by the National Institute for Health Research (NIHR) Cambridge Biomedical Research Centre (IS-BRC-1215- 20014). The NIHR Cambridge Biomedical Research Centre is a partnership between Cambridge University Hospitals NHS Foundation Trust and the University of Cambridge, funded by the NIHR. The views expressed are those of the authors and not necessarily those of the NHS, the NIHR or the Department of Health and Social Care. LJC has an advisory role at the NIHR BRC Nutritional Biomarker Laboratory.

We thank Tabasum Tabasum (NIHR BRC Nutritional Biomarker Laboratory) who performed the sample stability assays and Dave Collins (NIHR BRC Dietary Assessment, Anthropometry and Physical Activity Platform, MRC Epidemiology Unit, University of Cambridge) for providing NDNS data. Some of the experimental work was performed at the MRC Elsie Widdowson Laboratory, Cambridge, UK which closed in December 2018.

Competing Interests

The authors have no competing interests to declare.

Protocol for processing of whole blood samples required to produce washed erythrocyte specimens for the erythrocyte transketolase activity coefficient (ETKAC) assay

Materials

- Refrigerated centrifuge with shielded rotor (see note 1)

# Source of deionised water

- Saline. Approximately 30 ml of a 0.9% saline solution is required per participant for erythrocyte washing

- Pastettes

- Freezer (not self-defrosting) (see note 2)

Sample processing procedure

      - Pre-chill centrifuge to 4 °C

      - Mix all blood containers by inverting slowly 10 times

      - Centrifuge the tubes at 4 °C according to manufacturer’s instructions or local protocols for collection of plasma

      - Remove the plasma according to local protocols

      - Remove buffy coat top layer from the tube using a Pastette

      - Re-suspend the erythrocytes in normal saline (0.9% w/v NaCl solution) with approximately 2 times the volume of the erythrocytes. Gently invert the tube 10 times to mix

      - Centrifuge the tubes at 4 °C and 2000 g for 10 minutes and discard the supernatant

      - Repeat steps 6 and 7 twice for a total of 3 washes or until the supernatant is clear

      - Completely remove the supernatant

      - Washed erythrocytes should be stored at -70 °C

Notes for processing whole blood samples

Note 1. If a refrigerated centrifuge is not available then pre-cooling of centrifuge buckets and blocks at +4 °C is recommended.

Note 2. Washed erythrocyte samples should be frozen at -70 °C. Where this is not possible, samples should be stored at maximum -20 °C and transferred to -70 °C or below as soon as possible.

Protocol for the measurement of erythrocyte transketolase activity coefficient (ETKAC)

Materials

Equipment

- Incubator cabinet set at 37 °C

- Thermo Multiskan FC plate reader with incubator and 340 nm filter (Thermo Fisher Scientific, UK. Catalogue Number: 51119100) (see note 1), programmed appropriately for this assay

- Single channel pipettes, suitable for delivering 50 µl and 30 µl with correct tips (see note 2)

- Multichannel pipettes suitable for delivering 15 µl, 200 µl and 500 µl with correct tips (see note 2)

- Refrigerated centrifuge with shielded rotor

- pH meter

- Orbital plate shaker with speed control

- Glassware: 500 ml & 1 l graduated glass bottles, 1 l glass measuring cylinder

# Consumables

# 50 ml Falcon tubes (Corning Optical Communications, Flintshire, UK Catalogue Number: CLS430829), or equivalent

# 1 or 2 ml Pastettes

# Troughs for multichannel pipettes

# Greiner UVstar 96-well plates (Greiner Bio-One Ltd, Gloucestershire, UK. Catalogue Number: 655801), or equivalent (must be flat-bottomed and UV transparent at 340 nm)

# Screw-cap 2 ml Sarstedt tubes (Satstedt Ltd., Leicester, UK Catalogue Number: 72.730.406 ) or equivalent

# Self-adhesive plate sealers

# Source of deionised water (recommended resistivity of > 18.2 MΩ.cm)

# Method

# Reagent preparation (Table P1)

# 2M Hydrochloric Acid

# Add 417 ml deionised water to a gradated glass container
In a fume cupboard, make up to 500 ml with 37% hydrochloric acid
Store at room temperature

# Tris buffer

# Dissolve 12.1 g Trizma base in 700 - 800 ml deionised water. Buffer to pH 7.6 with 2M hydrochloric acid; make up to 1 l, mix by inversion
Store at +4 °C for up to one month

# 10% Tween 20 (see note 3)

# Measure 9 ml deionised water into a graduated tube
Add 1 ml Tween 20 below the meniscus using a Pastette, using the water meniscus and graduations to determine the volume (see note 4)
Cap the tube and mix by inversion
Store at +4 °C for up to 6 months

# 1% Tween 20

# Add 1 ml of 10% Tween 20 into a graduated tube using a pipette 
Make up to 10 ml with deionised water
Cap the tube and mix by inversion
Store at +4 °C for up to 6 months

# Thiamine pyrophosphate (TPP) (thiamine diphosphate, ThDP)

# Prepare a solution of 3.6 mg/ml TPP in Tris buffer. 1 ml is needed per plate, approximately 0.2 – 0.5 ml pipetting excess recommended
Weigh between 4.5 and 5.5 mg of TPP and calculate the volume of Tris buffer required; for example: 5 mg TPP / 1.39 ml Tris buffer
Prepare on day of analysis
Store at +4 °C

# Ribose-5-phosphate (R5P)

# Prepare a solution of 82.5 mg/ml R5P in Tris buffer
1 ml is needed per plate, approximately 0.2 – 0.5 ml pipetting excess recommended
Weigh between 100 and 120 mg of R5P and calculate the volume of Tris buffer required; for example: 100 mg R5P / 1.21 ml Tris buffer
Prepare on day of analysis
Store at +4 °C

# Nicotinamide adenine dinucleotide (NADH)

# Prepare a solution of 3.75 mg/ml NADH in Tris buffer 
1 ml is needed per plate, approximately 0.2 – 0.5 ml pipetting excess recommended
Weigh between 4.5 and 6 mg of NADH and calculate the volume of Tris buffer required; for example: 5 mg NADH / 1.33 ml Tris buffer
Prepare on day of analysis
Store at +4 °C

# αGDH-TPI

# 34 units required for each plate
Units per bottle vary from lot-to-lot; calculate volume required for each bottle of enzyme; for example, if the bottle contains 5000 U/4.8 ml, for 34 U use 32 µl (34 x 4.8/5000 = 32). 
Added directly to working reagent, as per instructions, on day of analysis
Store at +4 °C

# Working reagent

# Make up on day of analysis per plate during 15-minute plate incubation
Add the following to a 50 ml Falcon tube using a pipette:

# 20 ml Tris buffer

# 1 ml ribose-5-phosphate at 82.5 mg/ml

# 1 ml NADH at 3.75 mg/ml

# 32 µl GDH-TPI (see above; depends on activity of purchased stock)

# 80 µl 1% Tween-20

# Invert ten times to mix
Store at +4 °C until use

# Procedure

# 20 samples plus 3 quality control (QC) samples can be analysed in duplicate on a 96-well microplate

# The identification numbers of the samples, QC and batch number should be recorded in a laboratory notebook, together with the mass of reagents as weighed, and incubation timings

  - Preparation of equipment

The following steps must be taken to prepare equipment and reagents

    - Turn on the incubator and allow it to reach 37 °C (see note 5)

    - Turn on the refrigerated centrifuge and set it to cool to 4 °C

    - Turn on the plate reader and associated computer. Set temperature to 37 °C and allow at least 20 minutes for temperature to equilibrate (see note 6)

    - Remove NADH, R5P and TPP from -20 °C storage and allow to equilibrate to room temperature

    - Remove the controls and samples out from -70 °C storage and allow them to equilibrate to room temperature

    - Label a corresponding Sarstedt 2 ml tube for each of the samples

  - Preparation of hemolysates

    - Once the erythrocyte samples are thawed, mix the samples thoroughly by inversion

    - To lyse the cells, mix 1 volume of washed erythrocytes with 2 volumes of deionised water (see note 7). Hemolysates may be stored at -70 °C prior to assay if required

    - Mix thoroughly and leave at room temperature for 10 minutes for lysis to occur (see note 8)

  - Analysis

    - Centrifuge the samples and controls in the refrigerated centrifuge for 10 minutes at 4000 g

    - Produce ten-fold dilution of sample and control hemolysates with Tris buffer; e.g. 50 µl plus 450 µl Tris buffer in a 2 ml Sarstedt tube. Invert ten times to mix

    - Pipette 30 µl quadruplicates of each sample and QC as per the plate map in Figure P1. The QC samples should be assayed at the positions shown, ensuring that the position of any individual QC is varied from assay to assay. Assay buffer is used for the four blanks

    - Add 15 µl Tris to rows A, B, E, F  (basal) and 15 µl TPP to rows C, D, G, H  (activated) using a multichannel pipette

    - Cover the plate with a self-adhesive sealer and mix the plate on the plate shaker at 600 RPM for 30 seconds

    - Carefully place the plate in the incubator at 37 °C for 15 minutes

    - Prepare working reagent (see step 1.9)

    - Using a multichannel pipette, add 200 µl of working reagent per well

    - Mix on the plate shaker at 600 RPM for 30 seconds (do not cover)

    - Inspect wells for bubbles, if one is present it is advisable to burst it with a clean micropipette tip

    - Read each well of the plate at 340 nm 60 times at 1 minute intervals, with reader set to briefly shake the plate for 5 seconds before each read

  - Calculation of results

The first 40 readings, where the temperature of the plate equilibrates with that of the reader, are not included in the calculations.

    - Calculate the Δ change in absorbance for each well, during the last 20 readings

The mean blank rate is calculated, and subtracted from every “basal” and “activated” rate. It is advisable to check the four estimates of “blank” and if one deviates significantly from the other three it should be discounted.

In the assay each sample is measured in duplicate with addition of TPP (activated) and in duplicate without addition of TPP (basal). Duplicate measurements are required for each to allow interpretation of reproducibility of the pipetting of samples and reagents. Before the ETKAC can be calculated the imprecision of the activated and basal duplicates needs to be assessed; ETKAC should not be calculated if imprecision of duplicate measurements exceeds 10% for either the basal or activated measurement. To calculate imprecision of duplicates use the following formula:

Δ change in absorbance duplicate 1 / Δ change in absorbance duplicate 2

A ratio of 0.9 – 1.1 is acceptable.

    - Calculate the mean of each activated and basal duplicate pair

For each sample, divide the mean of the activated duplicate wells by the mean of the basal duplicate wells to calculate the ETKAC.

  - Quality control procedures

It is recommended that the kinetic trace of the readings are checked for anomalies, such as spikes (which indicates a plate reader malfunction, a bubble or electrical interference) or tapering of the reaction (which indicates insufficiency of reagents).

Samples with basal and/or activated duplicate ratios outside range of 0.9 - 1.1 should be rerun. If poor duplicate ratios occur regularly, this can be due to either a user error (e.g. poor pipetting technique) or mechanical error (e.g. plate reader malfunction).

Internal QC material

It is recommended that 3 separate internal control samples, as well as a reagent blank are run with each batch. The mean ETKAC value and the delta change in absorbance, for both the activated and basal duplicates should be recorded for each level of QC and the reagent blank. The rolling mean, standard deviation and %CV should be calculated for each recorded parameter. After 10 batches have been run the mean +/- 2 standard deviations of ETKAC values can be used as acceptance limits. The other parameters can be used to aid in interpretation of “failed” batches.

To produce internal QC material, washed erythrocytes from single donors should be lysed in bulk following the protocol above for preparation of hemolysates. The hemolysate should then be aliquoted and stored at -70 °C for analysis with each batch. It is important not to mix donations from different donors because clotting is likely to occur, as in a transfusion reaction.

QC ‘A’: Ideally produced from blood drawn from a thiamine deficient donor. If it is not possible to identify a thiamine deficient donor, a donor with an ETKAC close to the cut-off should be selected.

QC ‘B’: Should be produced from blood drawn from a donor without thiamine deficiency

QC ‘C’: “Low Activity” can be derived from any single-donor human washed erythrocyte sample, diluted 1:1 with physiological saline. This is to assess the robustness of the assay in samples with lower concentration of enzyme (for example in basal samples in thiamine deficiency).

Notes for the measurement of ETKAC

Note 1. To achieve acceptable analytical imprecision, it is crucial that the incubation temperature (37 °C) is maintained evenly across the plate during the read stage; the activity of the enzymes used for these reactions is temperature-dependent. The Thermo Multiskan plate reader is specified because of the design of the heating elements; unusually, the chamber is heated by full-sized plates above and below the microplate. Most microplate readers have chambers heated from one corner and cannot be used for this assay. The protocol described in Appendix A can be followed to assess the imprecision of a plate reader specifically for this assay and it is recommended that this protocol is followed before commencing ETKAC analysis and then yearly, or when troubleshooting. The protocol in Appendix A can also be used when assessing other plate readers for suitability with this assay.

Note 2. Any pipettes which conform to ISO8655 can be used.

Note 3. In the assay, absorbance is read repeatedly down through the well; each time the plate reader lands on one column of wells the position will be slightly different. Therefore, in order to have a constant path length it is necessary to ensure that the meniscus is flat, and this is achieved with the use of Tween 20.

Note 4. The use of a Pastette is recommended rather than a pipette due to the viscosity of Tween 20.

Note 5. The specifications of the plate reader state that the room temperature must be at least 5 C° below the set temperature and laboratories may require cooling air conditioning to achieve this.

Note 6. It is recommended that the plate reader is programmed to run the protocol before commencing analysis.

Note 7. For example, 200 µl washed erythrocytes and 400 µl of deionised water. The minimum recommended volume is 20 µl washed erythrocytes and 40 µl of deionised water.

Note 8. The hemolysate can be stored at -70 °C post-analysis for repeat batches or other assays.

Appendix A – Protocol to assess the suitability of plate reader performance for the ETKAC assay

The purpose of this procedure is to determine the level of within-plate imprecision due solely to the variation of temperature within the incubation chamber of the plate reader.  It is not possible to measure this temperature directly in each well; the plate reader recorded “37 °C” applies only to a single site within the chamber. In this procedure the normal protocol is followed except only one sample is used and processed in bulk before transferring to 96-well plate for reading. This means that any variation in rate between wells results only from temperature differences between the wells.

It is recommended that this procedure is run at least once a year and after any engineer service visit, as the assay will require a plate reader to exceed the manufacturer specifications. It may also be useful to follow this procedure when troubleshooting. Routine inspections and checks of the plate reader by an approved service engineer are also recommended.

A1. Procedure

- Turn on the incubator and allow it to reach 37 °C

- Turn on the plate reader and associated computer. Set temperature to 37 °C and allow at least 20 minutes for temperature to equilibrate

- Remove sufficient vials of a QC material to provide 300 µl hemolysate and allow to equilibrate to room temperature on a roller mixer

- Mix thoroughly by inversion. Centrifuge the samples and controls in the refrigerated centrifuge for 10 minutes at 4000 g

- Pipette 300 µl QC material hemolysate and 2.7 ml Tris buffer into a conical flask  or Falcon tube

- Add 1.5 ml TPP (to ensure higher rates) and swirl gently to mix

- Incubate for 15 minutes at 37 °C

- Prepare working reagent (see step 1.9 of “Protocol for the measurement of erythrocyte transketolase activity coefficient (ETKAC)”)

- Add 19.2 ml working reagent and swirl gently to mix

- Pipette 200 µl of the mixture into each well of the 96-well plate

- Mix on the plate shaker at 600 RPM for 30 seconds (do not cover)

- Inspect wells for bubbles, if one is present it is advisable to burst it with a clean micropipette tip

- Read each well of the plate at 340 nm 60 times at 1 minute intervals, with reader set to briefly shake the plate for 5 seconds before each read

A2. Calculation of results

- The first 40 readings, where the temperature of the plate is being brought to that of the reader, are not included in the calculations

- Calculate the Δ change in absorbance at 340 nm for the last 20 readings for each well of the plate

- Inspect traces for “spikes” if any individual rate is aberrant. Wells in which the rate has been affected by a spike or non-linearity or from mis-pipetting should be excluded from the calculations

- Calculate the mean, standard deviation and %CV of the  Δ change in absorbance measurements for each row, each column and for the whole plate

- Inspect the results to look for trends in rate (i.e. in temperature) and also to look for edge effects (centre warmer than edges)

- If the CVs of the rows or columns demonstrate unevenness in the enzyme rate and thus in the temperature in the wells, or if there is an “edge effect”, the part of the plate affected is not suitable for this assay. The recommendation is that for rows and columns with the plate the CV should be < 2% and for the plate overall, < 3%

- Using notional “blank” rates (e.g. average blank rates obtained from QC data), calculate “ETKAC” for each set of four wells in the same way as would be done for an assay. Ideally each calculated ETKAC will be 1.00 because the same reaction is occurring in every well. Inspect the results for deviations from 1.00; this provides an indication of the within-batch imprecision to be expected from an assay

References

1.	Frank L.L. 2015. Thiamin in Clinical Practice. JPEN. J. Parenter. Enteral Nutr. 39: 503–520.

2.	Whitfield K.C., M.W. Bourassa, B. Adamolekun, et al. 2018. Thiamine deficiency disorders: diagnosis, prevalence, and a roadmap for global  control programs. Ann. N. Y. Acad. Sci. 1430: 3–43.

3.	Hiffler L., B. Rakotoambinina, N. Lafferty, et al. 2016. Thiamine Deficiency in Tropical Pediatrics: New Insights into a Neglected but Vital  Metabolic Challenge. Front. Nutr. 3: 16.

4.	McCann A., Ø. Midttun, K.C. Whitfield, et al. 2017. Comparable Performance Characteristics of Plasma Thiamine and Erythrocyte Thiamine  Diphosphate in Response to Thiamine Fortification in Rural Cambodian Women. Nutrients 9:.

5.	Gibson R.S. 2005. “Principles of Nutritional Assessment.,” Second New York: Oxford University Press.

6.	Collie J.T.B., R.F. Greaves, O.A.H. Jones, et al. 2017. Vitamin B1 in critically ill patients: needs and challenges. Clin. Chem. Lab. Med. 55: 1652–1668.

7.	Brin M. 1962. Effects of thiamine deficiency and of oxythiamine on rat tissue transketolase. J. Nutr. 78: 179–183.

8.	Smeets E.H., H. Muller & J. de Wael. 1971. A NADH-dependent transketolase assay in erythrocyte hemolysates. Clin. Chim. Acta. 33: 379–386.

9.	Mount J.N., E. Heduan, C. Herd, et al. 1987. Adaptation of coenzyme stimulation assays for the nutritional assessment of vitamins  B1, B2 and B6 using the Cobas Bio centrifugal analyser. Ann. Clin. Biochem. 24 ( Pt 1): 41–46.

10.	Vuilleumier J.P., H.E. Keller & E. Keck. 1990. Clinical chemical methods for the routine assessment of the vitamin status in human  populations. Part III: The apoenzyme stimulation tests for vitamin B1, B2 and B6 adapted to the Cobas-Bio analyzer. Int. J. Vitam. Nutr. Res. Int. Zeitschrift  fur Vitamin- und Ernahrungsforschung. J. Int. Vitaminol. Nutr. 60: 126–135.

11.	Public Health England. 2014. “National Diet and Nutrition Survey Results from Years 1, 2, 3 and 4 (combined) of the Rolling Programme (2008/2009 – 2011/2012).”

12.	Takeuchi T., E.H. Jung, K. Nishino, et al. 1990. The relationship between the thiamin pyrophosphate effect and the saturation status  of the transketolase with its coenzyme in human erythrocytes. Int. J. Vitam. Nutr. Res. Int. Zeitschrift  fur Vitamin- und Ernahrungsforschung. J. Int. Vitaminol. Nutr. 60: 112–120.

13.	EFSA Panel on Dietetic Products N. and A. (NDA), D. Turck, J.-L. Bresson, et al. 2016. Dietary reference values for thiamin. EFSA J. 14: e04653.

14.	Medicine I. of. 1998. “Dietary Reference Intakes for Thiamin, Riboflavin, Niacin, Vitamin B 6, Folate, Vitamin B 12, Pantothenic Acid, Biotin, and Choline.” Washington (DC).

15.	Shaw N.-S., J.-L. Wang, W.-H. Pan, et al. 2007. Thiamin and riboflavin status of Taiwanese elementary schoolchildren. Asia Pac. J. Clin. Nutr. 16 Suppl 2: 564–571.

16.	Bates C.J. 1997. Vitamin Analysis. Ann. Clin. Biochem. 34: 599–626.

17.	Wood B., K.J. Breen & D.G. Penington. 1977. Thiamine status in alcoholism. Aust. N. Z. J. Med. 7: 475–484.

18.	Soukaloun D., S.J. Lee, K. Chamberlain, et al. 2011. Erythrocyte Transketolase Activity, Markers of Cardiac Dysfunction and the Diagnosis of Infantile Beriberi. PLoS Negl. Trop. Dis. 5: e971.

19.	Taylor A.J., D. Talwar, S.J. Lee, et al. 2020. Comparison of Thiamin Diphosphate High-Performance Liquid Chromatography and  Erythrocyte Transketolase Assays for Evaluating Thiamin Status in Malaria Patients without Beriberi. Am. J. Trop. Med. Hyg.

20.	Buttery J.E. & P.R. Pannall. 1985. Kinetic transketolase assay: use of whole-blood hemolysate as the sample. Clin. Chem. 31: 1086.

21.	Cheng C.H., M. Koch & R.E. Shank. 1976. Leukocyte transketolase activity as an indicator of thiamin nutriture in rats. J. Nutr. 106: 1678–1685.

22.	Leukocyte transketolase activity: an indicator of thiamin nutriture. 1977. Nutr. Rev. 35: 185–187.

23.	Ali M., C.J. Gubler, J. Al Saleh, et al. 1987. A comparison of transketolase assay and transketolase and lactate dehydrogenase  activity levels in whole blood and red cell hemolysates and in leukocytes. Comp. Biochem. Physiol. B. 87: 833–835.

24.	Anderson S.H. & A.D. Nicol. 1985. Kinetic transketolase assay: use of whole-blood hemolysate as the sample. Clin. Chem. 31: 2044.

25.	Anderson S.H. & A.D. Nicol. 1986. A fluorimetric method for measurement of erythrocyte transketolase activity. Ann. Clin. Biochem. 23 ( Pt 2): 180–189.

26.	Puxty J.A., A.E. Haskew, J.G. Ratcliffe, et al. 1985. Changes in erythrocyte transketolase activity and the thiamine pyrophosphate effect  during storage of blood. Ann. Clin. Biochem. 22 ( Pt 4): 423–427.

27.	Bayoumi R.A. & S.B. Rosalki. 1976. Evaluation of methods of coenzyme activation of erythrocyte enzymes for detection of  deficiency of vitamins B1, B2, and B6. Clin. Chem. 22: 327–335.

28.	van Dokkum W., J. Schrijver & J.A. Wesstra. 1990. Variability in man of the levels of some indices of nutritional status over a 60-d  period on a constant diet. Eur. J. Clin. Nutr. 44: 665–674.

29.	Public Health England. 2018. “National Diet and Nutrition Survey. Results from Years 7 and 8 (combined) of the Rolling Programme (2014/15 to 2015/16).”

30.	Kaczmarek M.J. & P.F. Nixon. 1983. Variants of transketolase from human erythrocytes. Clin. Chim. Acta. 130: 349–356.

31.	Nixon P.F., M.J. Kaczmarek, J. Tate, et al. 1984. An erythrocyte transketolase isoenzyme pattern associated with the  Wernicke-Korsakoff syndrome. Eur. J. Clin. Invest. 14: 278–281.

32.	Kaufmann A., S. Uhlhaas, W. Friedl, et al. 1987. Human erythrocyte transketolase: no evidence for variants. Clin. Chim. Acta. 162: 215–219.

33.	Tate J.R. & P.F. Nixon. 1987. Measurement of Michaelis constant for human erythrocyte transketolase and thiamin  diphosphate. Anal. Biochem. 160: 78–87.

34.	Martin P.R., B.A. McCool & C.K. Singleton. 1995. Molecular genetics of transketolase in the pathogenesis of the Wernicke-Korsakoff  syndrome. Metab. Brain Dis. 10: 45–55.

35.	Blansjaar B.A., R. Zwang & B.G. Blijenberg. 1991. No transketolase abnormalities in Wernicke-Korsakoff patients. J. Neurol. Sci. 106: 88–90.

36.	Kochetov G.A. & O.N. Solovjeva. 2014. Structure and functioning mechanism of transketolase. Biochim. Biophys. Acta 1844: 1608–1618.

37.	Huang Y., R.A. Gibson & T.J. Green. 2020. Measuring thiamine status in dried blood spots. Clin. Chim. Acta 509: 52–59.

38.	DREYFUS P.M. 1962. Clinical application of blood transketolase determinations. N. Engl. J. Med. 267: 596–598.

39.	Pekovich S.R., P.R. Martin & C.K. Singleton. 1996. Thiamine pyrophosphate-requiring enzymes are altered during pyrithiamine-induced  thiamine deficiency in cultured human lymphoblasts. J. Nutr. 126: 1791–1798.

40.	Price J., A.E. Clague, R.A. Kerr, et al. 1991. In thiamine deficiency, activation of erythrocyte transketolase by thiamine in vivo  exceeds activation by cofactor in vitro. Clin. Chim. Acta. 202: 39–45.

41.	Spooner R.J., R.A. Percy & A.G. Rumley. 1979. The effect of erythrocyte ageing on some vitamin and mineral dependent enzymes. Clin. Biochem. 12: 289–290.

42.	Talwar D., H. Davidson, J. Cooney, et al. 2000. Vitamin B(1) status assessed by direct measurement of thiamin pyrophosphate in  erythrocytes or whole blood by HPLC: comparison with erythrocyte transketolase activation assay. Clin. Chem. 46: 704–710.

43.	Howard J.M. 2000. Assessment of vitamin B(1) status. Clin. Chem. 46: 1867–1868.

Figure Legends

Figure 1. The enzymatic reactions in the ETK activity assay. The assay monitors the decrease in absorbance at 340 nm related to NADH oxidation.

Figure 2. Deming regression of ETKAC measured in erythrocytes hemolysates collected in either ETDA or lithium heparin (LH) blood tubes from 15 adults. Open circles, observed data points; solid line, regression line; dashed line, line of equality. Slope (95% CI)  = 0.88 (0.37, 1.39), intercept (95% CI) = 0.26 (-0.44, 0.68).

Figure 3. Deming regression of ETKAC measured in erythrocyte hemolysates prepared from LH and EDTA whole blood collected from 15 adults and processed with 2 h of collection or after refrigeration for 24 h. Dashed line is the line of equality.

Figure 4. Deming regression of repeat analysis of erythrocyte hemolysates after storage for 6 months at -70 °C. Open circles, observed data points; solid line, regression line; dashed line, line of equality. Slope (95% CI)  = 1.06 (0.92, 1.20), intercept (95% CI) = -0.07 (-0.24, 0.09) (n=59).

Figure P1. Plate map for ETKAC analysis. Sample and QC (30 µl) are added to each of four wells on the plate. To the shaded rows A, B, E, F (basal wells) is added 15 µl Tris buffer and to rows C, D, G, H (activated wells) 15 µl TPP is added. Abbreviations: BLK, reagent blank (Tris buffer); QC, internal quality control sample; 1-20, samples.