The Artificial Pancreas

Charlotte K Boughton and Roman Hovorka

Wellcome Trust-Medical Research Council Institute of Metabolic Science,

University of Cambridge, Cambridge, UK

Corresponding author

Charlotte K Boughton MRCP PhD

University of Cambridge Metabolic Research Laboratories and

NIHR Cambridge Biomedical Research Centre,

Wellcome Trust-MRC Institute of Metabolic Science, Box 289,

Addenbrooke’s Hospital, Hills Road, Cambridge CB2 0QQ, UK

tel: +44 1223 769 066, , e-mail: cb2000@medschl.cam.ac.uk

Disclosure of funding: Supported by the National Institute of Health Research Cambridge Biomedical Research Centre, Efficacy and Mechanism Evaluation National Institute for Health Research, The Leona M. & Harry B. Helmsley Charitable Trust, JDRF, National Institute of Diabetes and Digestive and Kidney Diseases and Diabetes UK.

Manuscript word count: 2628

Abstract word count: 168

Visuals:

Figures: 3

Abstract

Purpose of review:

Advances in diabetes technologies have enabled the development of artificial pancreas (closed-loop) systems for people with diabetes. We review the key studies which have led to the adoption of the artificial pancreas in clinical practice and consider ongoing challenges and areas for future enhancements.

Recent findings:

Studies have demonstrated safety and efficacy of closed-loop insulin delivery systems in free-living settings over periods of up to 6 months for children and adults with type 1 diabetes. Since 2017, four hybrid closed-loop systems have been approved by regulatory bodies worldwide, but these systems are not entirely automated, requiring user interaction to deliver mealtime insulin boluses. Improving usability of these devices in the real-world setting is an important challenge.

Summary:

The artificial pancreas has become the gold standard for the treatment of type 1 diabetes. First generation systems are increasingly being adopted in clinical practice, however further work is required, developing advanced systems and faster acting insulin analogues to allow complete automation and further reduce the burden of type 1 diabetes.

Keywords:

- Artificial pancreas

- Hybrid closed-loop

- Automated insulin delivery

- Type 1 diabetes

- Inpatient diabetes

- Usability

Introduction

Pancreas and islet transplantation are considered, at present, last resource approaches to the treatment of type 1 diabetes (T1D), suitable only for those with significant burden of hypoglycaemia where other approaches have failed. Due to the nature and risks of the associated immunosuppressive regimen, such treatments are not practicable for children and young people. Physiological replacement of endocrine pancreatic function using exogenous insulin remains the mainstay of treatment for people with T1D. Recent years have seen significant advances in the development of artificial pancreas (closed-loop) systems which combine real-time sensor glucose data with an algorithm to direct insulin delivery. In this review we will consider recent studies demonstrating safety and efficacy of the artificial pancreas and discuss future advances required to more closely mimic glucose control by endogenous pancreatic function.

Background

Maintaining glycaemic control is pivotal to diabetes management. Attainment of intensive glycaemic control was associated with significantly decreased rates of development and progression of microvascular complications compared with standard glycaemic control in people with T1D in the Diabetes Control and Complications Trial (1).

The healthy pancreatic beta-cell is highly sensitive to changes in glucose concentrations, rapidly releasing insulin in a glucose-dependent manner. In T1D, where the beta-cells are dysfunctional, individuals are required to frequently monitor glucose levels and manually adjust insulin doses to optimise glucose control. Technological advances have provided significant benefits for diabetes management. Introduction of continuous insulin infusion pumps in the 1970s allowed more physiological insulin delivery compared with injections. Intermittent capillary blood glucose testing is being superseded by continuous glucose monitors (CGM) which permit real-time monitoring of interstitial fluid glucose levels. Despite increased adoption of these technologies, only a minority of children and adults meet recommended glycaemic targets and the burden of managing T1D remains significant. (2)

Automated insulin delivery systems, while initially conceptualised in the 1960s, only became feasible after CGM was shown to be accurate and reliable. Integration of CGM and pumps (sensor augmented pump therapy) allows for the simplest form of automated insulin delivery with suspension of insulin at a threshold glucose level (low glucose suspend) or when a threshold glucose level is predicted to be reached (predictive low glucose suspend). Closed-loop systems are more sophisticated, combining information provided by CGM with an algorithm, which autonomously calculates and directs the insulin pump to deliver insulin based on real-time glucose trends (Figure 1). Dual-hormone closed-loop systems which control both insulin and glucagon or pramlintide are also in development.

Regulatory approval of closed-loop systems

Approval of the first closed-loop system (Medtronic MiniMed 670G) by the Food and Drug Administration in 2017 represented a milestone in translation of prototype closed-loop systems into clinical practice (Figure 2). This hybrid closed-loop system was shown in the pivotal study to be safe and effective in people aged ≥14 years with T1D. (3) Approval in Europe was received in 2018 and the licence was recently extended to children aged ≥7 years (3). The term hybrid closed-loop reflects the need for user input as delays in subcutaneous insulin absorption mean user-initiated mealtime insulin delivery (bolusing) is required. The ultimate goal is for complete automation. Three other hybrid closed-loop systems have since received regulatory approval; Tandem Control-IQ in the US and Diabeloop and CamAPS FX in Europe.

Meta-analysis of efficacy and safety of closed-loop systems

Closed-loop system outcomes are reported in terms of efficacy measured as time spent with sensor glucose levels within the target range (70-180 mg/dL; 3.9-10.0 mmol/L in non-pregnant individuals) and safety outcomes. Increasing attention is being given to usability outcomes. A recent meta-analysis examined 40 studies: 31 used a single-hormone system and 9 used a dual-hormone approach (4). Artificial pancreas systems outperform current standard of care (sensor-augmented pump therapy) across a number of glycaemic metrics including an increase in time in target glucose range of 9.6 percentage points (equating to over two and a half additional hours in target/day) and a reduction in time spent in hypoglycaemia of 1.5 percentage points (22 minutes/day). Closed-loop system use is associated with reduction in HbA1c of 0.26%.

Key outpatient artificial pancreas studies

Assessment of the efficacy of closed-loop insulin delivery from large clinical trials in free living unsupervised settings in representative populations is crucial critical to support adoption and reimbursement for people with T1D.

Diabeloop DBLG1 closed-loop system

A large multicentre randomised crossover study in France compared the Diabeloop DBLG1 hybrid closed-loop system with sensor-augmented pump therapy for 12 weeks in 68 adult insulin pump users with T1D in the home setting but with remote monitoring (5). Baseline HbA1c was 7.6% (59.4 mmol/mol). The proportion of time with sensor glucose in the target range was 9.2 percentage points greater with hybrid closed-loop than with control therapy (69% vs 59%). The time spent in hypoglycaemia (<70 mg/dL; 3.9 mmol/L) was significantly lower (30 minutes less per day) with closed-loop than during the control period. Mean HbA1c was reduced by 0.29% with closed-loop compared with 0.14% in the control group without reaching statistical significance for between group comparison. During the study, five severe hypoglycaemic episodes occurred in the closed-loop group and three in the sensor-augmented pump therapy group; these events were attributed to pump hardware malfunctions or human error. The DBLG1 hybrid closed-loop system has received CE mark for use in adults with T1D and commercialisation is in progress.

Cambridge closed-loop system

A multicentre study across UK and US compared hybrid closed-loop with sensor-augmented pump therapy in a diverse population of 86 children, adolescents and adults with sub-optimal glycaemic control despite pump therapy (baseline HbA1c 8.2% (67 mmol/mol at recruitment) over a period of 12 weeks of unsupervised free-living (6). The time spent with glucose in target range was 10.8 percentage points higher with closed-loop compared with control therapy (65% vs 54%). Time spent in hypoglycaemia (<70 mg/dL; 3.9 mmol/L) was also significantly reduced with closed-loop. HbA1c improved in both groups during the run-in period but the closed-loop group had a further reduction from 8.0% (64 mmol/mol) after run-in to 7.4% (57 mmol/mol) after the 12-week intervention period, while in the control group, the HbA1c was 7.8% (62 mmol/mol) after run-in and 7.7% (61 mmol/mol) after the intervention period. After adjusting for baseline HbA1c, the reduction in HbA1c was 0.36% (4.0 mmol/mol) greater in the closed-loop group. The Cambridge CamAPS FX hybrid closed-loop system received CE mark in March 2020 for use in children aged 1 year and older and adults with T1D and is soon to be commercially available.

Tandem Control-IQ closed-loop system

The largest and longest closed-loop study to date was a six month parallel design randomised controlled trial including 168 adolescents (≥14 years) and adults (7). Participants included both pump users and those using multiple daily insulin injections and were assigned in a 2:1 ratio to receive treatment with a closed-loop system or sensor-augmented pump therapy. Baseline HbA1c was 7.4% (57.4 mmol/mol). The percentage of time glucose was within target range increased in the closed-loop group from 61% at baseline to 71% during the 6 months and remained unchanged at 59% in the control group; an 11 percentage points difference equating to an additional two and a half hours/day (Figure 3). The percentage of time that the glucose level was <70 mg/dl (3.9 mmol/L) was lower with closed-loop than in the control group. The difference in HbA1c after 6 months was 0.33 percentage points lower with closed-loop. Closed-loop usage was very high; 90% over the 6 month intervention period and all 168 participants completed the study suggesting high usability within the clinical trial setting. The Tandem Diabetes Care Control-IQ hybrid closed-loop system received FDA approval in December 2019 for use in people aged ≥14 years with T1D and is commercially available in the US.

The artificial pancreas in very young children

Diabetes management in very young children is particularly challenging due to unpredictable eating habits and activity. A secondary analysis of closed-loop studies showed higher variability of insulin requirements in younger children, particularly overnight, compared with older children and adults, identifying a clear need for automated glucose-responsive systems in this population.(8) Safety and efficacy of hybrid closed-loop therapy in this population was demonstrated in a multicentre randomised crossover study which compared hybrid closed-loop using diluted insulin with hybrid closed-loop using standard strength insulin over a three week free-living period. (9) While there was no difference in the time spent in target glucose range or time in hypoglycaemia between the two interventions, overall the time spent in target glucose range was over 70% with little time in hypoglycaemia. Important quality of life improvements with closed-loop therapy were reported for both very young children and caregivers (10).

The artificial pancreas in pregnant women with T1D

Glycaemic targets are much tighter during pregnancy as maternal hyperglycaemia is associated with increased rates of stillbirth, neonatal death, preterm delivery and macrosomia. Insulin requirements typically increase 2-3 fold throughout pregnancy, with high day-to-day variability. Despite intensification of insulin therapy, few women meet glycaemic targets and the burden of hypoglycaemia is high (11). In a randomized crossover study, hybrid closed-loop therapy was safe and effective in pregnant women with T1D (12). Closed-loop insulin delivery overnight was associated with increased time spent in the tighter target glucose range for pregnancy (63-140 mg/dL; 3.5-7.8 mmol/L) compared with sensor augmented pump therapy (75% vs. 60%), without increasing hypoglycaemia. During an extension phase of day-and-night closed-loop insulin delivery (including antenatal hospitalisation, labour and delivery), glucose levels were in target range 69% of the time. The impact of closed-loop therapy on fetal and maternal perinatal outcomes remains to be determined but larger outcome studies are underway.

Dual-hormone artificial pancreas

First-generation closed-loop systems only automate insulin delivery. In T1D in addition to insulin deficiency, glucagon release in response to hypoglycaemia is also affected and amylin which is co-secreted with insulin is deficient. Replicating islet biology mechanically will likely require replacement and restoration of more than just insulin. Several dual-hormone closed-loop systems are being studied incorporating glucagon and/or amylin to improve glycaemic control beyond delivery of insulin alone. Addition of a second hormone has challenges including developing dual-chambered infusion pumps and infusion sets, formulating stable glucagon for pump delivery, and increased complexity in control algorithms.

Addition of glucagon theoretically confers additional protection from hypoglycaemia and also permits more aggressive insulin delivery. In outpatient trials comparing single-hormone (insulin alone) and dual-hormone (insulin and glucagon) systems, addition of glucagon reduced the time spent in hypoglycaemia by 33 minutes/day without impacting on the time spent in target glucose range (13). An outpatient crossover study compared a dual-hormone (insulin and glucagon) system with usual care over 11 days at home (14); there was a reduction of 18 minutes/day spent with glucose levels <60 mg/dL (3.3 mmol/L) compared to the control group. The time spent in target glucose range (70-180 mg/dL; 3.9-10 mmol/L) was superior with the dual-hormone system compared to usual care (79% vs 62%) equating to an additional four hours in target each day. Nausea scores were greater with the dual-hormone system.

Pramlintide has favourable actions on delaying gastric emptying and reducing the magnitude of post-prandial glucose excursions. A randomised crossover trial compared a rapid insulin-alone artificial pancreas with rapid insulin-and-pramlintide artificial pancreas systems in adults over a 24 hour inpatient stay (15). Compared with the rapid insulin-alone system, the rapid insulin-and-pramlintide system increased the time in range from 74% to 84%. The increased time in range observed was due to improved daytime control, typically a more challenging period for closed-loop glucose control (daytime in range increased from 63% to 78%). There were more gastrointestinal symptoms reported with the rapid insulin-and-pramlintide system but none of the symptoms were severe.

Additional longer studies with multi-hormone systems are required to demonstrate reductions in hypo- and hyperglycemia exposure compared with single-hormone systems. No dual-hormone systems have yet been approved for use in clinical practice.

Usability of the artificial pancreas

With increasing use of hybrid closed-loop systems in the real-world there has been a shift in focus to usability. First generation systems offer limited flexibility and require significant workload to maintain closed-loop operation (so called Auto Mode). A real-world study in the US reported that during the first year of hybrid closed-loop therapy with the Medtronic 670G system, 33% of users discontinued use (16). Reasons for discontinuation included CGM issues (calibrations), number of alarms, and too much time and workload required to make the system operate in Auto Mode (17).

Future systems must consider user interaction in the design to ensure the glycaemic benefits of hybrid closed-loop therapy are realised in the real-world setting. Flexibility and additional functions within the closed-loop system to allow for customisation are likely to be more appealing to users. Interoperability of the individual closed-loop devices (insulin pump, CGM and algorithm) which would allow users increased choice of system components to suit their needs, are being encouraged and are likely to improve usability and therefore outcomes.

Do-It-Yourself (DIY) artificial pancreas

Delays in closed-loop systems receiving regulatory approval led to communities including OpenAPS (Open Artificial Pancreas System), Loop, and AndroidAPS developing alternative non-commercial systems (18). Individuals build their own closed-loop system from commercially available, although sometimes out of warranty, insulin pumps, CGM devices, and open source algorithms. The lack of regulatory approval for these systems has the advantage of expeditious innovation cycles and increased functionality for customisation. Reports suggest these systems have several thousand users to date worldwide. Initial outcomes reported by this community are positive, but no randomized controlled trials have been reported. Healthcare professional responsibilities in supporting users of these unregulated systems remain controversial, and the FDA released a safety communication warning against using DIY systems.

Inpatient use of the artificial pancreas

Diabetes management in the hospital can be very challenging and dysglycaemia is associated with increased morbidity and mortality (19). A randomised controlled trial of adult inpatients with hyperglycaemia compared fully closed-loop insulin delivery (without meal time boluses) with standard insulin therapy (20). Closed-loop therapy was associated with almost six additional hours/day spent with glucose in the target range (100-180 mg/dL; 5.6-10 mmol/L) compared with standard insulin therapy (66 vs 42%). This was achieved without an increase in hypoglycemia. Similar superior glucose control with closed-loop (additional eight hours/day with glucose in target range) was observed in a smaller randomised controlled trial involving inpatients receiving nutrition support (enteral/parenteral nutrition) (21). Further studies are required to establish if improvements in glucose control with closed-loop can translate into improved patient outcomes, reduced staff workload and healthcare costs.

Limitations of the artificial pancreas

Despite considerable progress, there remain important limitations with current closed-loop systems. Delays in subcutaneous insulin absorption limit the possibility of fully automated systems for people with T1D and therefore user interaction is required for mealtime boluses to minimise post-prandial hyperglycaemia. Subcutaneous insulin infusion site issues can cause inconsistent insulin delivery which is problematic for closed-loop systems. Intraperitoneal insulin infusion would reduce lag times associated with subcutaneous insulin delivery and may improve glucose control.

Current systems require two separate devices attached to the body (insulin pump and CGM) and attempts to reduce the size of these devices or combine them are in development and will improve device burden and wearability particularly for young children. Although closed-loop systems might be expected to reduce night-time anxiety and improve sleep quality, there are reports of negative impact of technology on sleep with increased alarm burden.

An artificial pancreas which incorporates inputs in addition to glucose concentration (e.g. heart rate) with more advanced algorithms may allow more physiological insulin delivery and therefore tighter glucose regulation.

Conclusion

Considerable advances have been made recently with translation of the artificial pancreas into clinical practice. Current closed-loop systems still require user input and are associated with device burden, however are safer and/or more effective than alternative approaches. While transplantation has the benefit of potential insulin independence, surgical and immunosuppression risks make this approach unsuitable for many. Immunotherapies and cell-based therapies which might prevent the decline or even restore endogenous insulin secretion remain in development phases, so at present technology is the mainstay of management of T1D.

Key points:

- The artificial pancreas (closed-loop system) has recently been translated into clinical practice for people with type 1 diabetes

- Hybrid closed-loop systems can achieve superior glucose control in people with type 1 diabetes

- Increasing attention is being given to usability to realise the glycaemic benefits of closed-loop therapy in real-world settings

- Closed-loop systems have limitations but at present are safer and/or more effective than alternative approaches to management of type 1 diabetes

Acknowledgements

Financial support and sponsorship.

Supported by the National Institute of Health Research Cambridge Biomedical Research Centre, Efficacy and Mechanism Evaluation National Institute for Health Research, The Leona M. & Harry B. Helmsley Charitable Trust, JDRF, National Institute of Diabetes and Digestive and Kidney Diseases and Diabetes UK.

Conflicts of interest.

RH reports having received speaker honoraria from Eli Lilly and Novo Nordisk, serving on advisory panel for Eli Lilly and Novo Nordisk, receiving license fees from B. Braun and Medtronic; having served as a consultant to B. Braun, patents and patent applications related to closed-loop, and being shareholder of CamDiab. CB declares no conflict of interest associated with this manuscript.

References

1.	Nathan DM, Genuth S, Lachin J, et al. The effect of intensive treatment of diabetes on the development and progression of long-term complications in insulin-dependent diabetes mellitus. New England Journal of Medicine. 1993;329(14):977-86.

2.	Foster NC, Beck RW, Miller KM, et al. State of Type 1 Diabetes management and outcomes from the T1D Exchange in 2016-2018. Diabetes Technology & Therapeutics. 2019;21(2):66-72.

3.	Garg SK, Weinzimer SA, Tamborlane WV, et al. Glucose outcomes with the in-home use of a hybrid closed-loop insulin delivery system in adolescents and adults with Type 1 Diabetes. Diabetes Technology & Therapeutics. 2017;19(3):155-63.

4.	Bekiari E, Kitsios K, Thabit H, et al. Artificial pancreas treatment for outpatients with type 1 diabetes: systematic review and meta-analysis. British Medical Journal. 2018;361.

5.	*Benhamou PY, Franc S, Reznik Y, et al. Closed-loop insulin delivery in adults with type 1 diabetes in real-life conditions: a 12-week multicentre, open-label randomised controlled crossover trial. The Lancet Digital Health. 2019;1(1):E17-E25.

- This study is a randomised crossover study comparing hybrid closed-loop therapy with sensor augmented pump therapy in adults. This is the first approved closed-loop system incorporating an insulin patch-pump.

6.	Tauschmann M, Thabit H, Bally L, et al. Closed-loop insulin delivery in suboptimally controlled type 1 diabetes: a multicentre, 12-week randomised trial. Lancet. 2018;392(10155):1321-9.

7.	**Brown SA, Kovatchev BP, Raghinaru D, et al. Six-month randomized, multicenter trial of closed-loop control in type 1 diabetes. New England Journal of Medicine. 2019;381(18):1707-17.

- This is the largest and longest study of a hybrid closed-loop system in adolescents and adults under free-living unsupervised conditions and included people using both insulin injections and pump therapy at baseline. The study had 100% retention and high closed-loop usage suggesting positive usability of this closed-loop system which has important implications for adoption and reimbursement.

8.	*Dovc K, Boughton C, Tauschmann M, et al. Young children have higher variability of insulin requirements: observations during hybrid closed-loop insulin delivery. Diabetes Care. 2019;42(7):1344-7.

- This secondary analysis of closed-loop studies reports higher variability of insulin requirements in very young children, particularly overnight, supporting expedited clinical practice adoption of closed-loop therapy in this population.

9.	*Tauschmann M, Allen JM, Nagl K, et al. Home Use of Day-and-Night Hybrid Closed-Loop Insulin Delivery in Very Young Children: A Multicenter, 3-Week, Randomized Trial. Diabetes Care. 2019;42(4):594-600.

- This pilot study is the first to report on the safety and efficacy of hybrid closed-loop therapy in very young children under free-living unsupervised conditions.

10.	Musolino G, Dovc K, Boughton CK, et al. Reduced burden of diabetes and improved quality of life: Experiences from unrestricted day-and-night hybrid closed-loop use in very young children with type 1 diabetes. Pediatric Diabetes. 2019;20(6):794-9.

11.	Murphy HR, Rayman G, Duffield K, et al. Changes in the glycemic profiles of women with type 1 and type 2 diabetes during pregnancy. Diabetes Care. 2007;30(11):2785-91.

12.	Stewart ZA, Wilinska ME, Hartnell S, et al. Closed-Loop Insulin Delivery during Pregnancy in Women with Type 1 Diabetes. New England Journal of Medicine. 2016;375(7):644-54.

13.	Haidar A, Messier V, Legault L, et al. Outpatient 60‐hour day‐and‐night glucose control with dual‐hormone artificial pancreas, single‐hormone artificial pancreas, or sensor‐augmented pump therapy in adults with type 1 diabetes: An open‐label, randomised, crossover, controlled trial. Diabetes, Obesity and Metabolism. 2017;19(5):713-20.

14.	El-Khatib FH, Balliro C, Hillard MA, et al. Home use of a bihormonal bionic pancreas versus insulin pump therapy in adults with type 1 diabetes: a multicentre randomised crossover trial. Lancet. 2017;389(10067):369-80.

15.	* Haidar A, Tsoukas MA, Bernier-Twardy S, et al. A Novel Dual-Hormone Insulin-and-Pramlintide Artificial Pancreas for Type 1 Diabetes: A Randomized Controlled Crossover Trial. Diabetes Care. 2020;43(3):597-606.

- This is the first study using a novel insulin-and-pramlintide artificial pancreas system and demonstrates benefits on glucose control (particularly in the daytime) compared with an insulin alone system.

16.	* Lal RA, Basina M, Maahs DM, et al. One Year Clinical Experience of the First Commercial Hybrid Closed-Loop System. Diabetes Care. 2019;42(12):2190-6.

- This study reports real-world outcomes with the first commercially available closed-loop system and highlights the importance of usability.

17.	Messer LH, Berget C, Vigers T, et al. Real world hybrid closed-loop discontinuation: Predictors and perceptions of youth discontinuing the 670G system in the first 6 months. Pediatric Diabetes. 2020;21(2):319-27.

18.	Lewis DM. Do-It-Yourself Artificial Pancreas System and the OpenAPS Movement. Endocrinology and Metabolism Clinics of North America. 2020;49(1):203-13.

19.	McDonnell ME, Umpierrez GE. Insulin therapy for the management of hyperglycemia in hospitalized patients. Endocrinology and Metabolism Clinics of North America. 2012;41(1):175-201.

20.	Bally L, Thabit H, Hartnell S, et al. Closed-loop insulin delivery for glycemic control in noncritical care. New England Journal of Medicine. 2018;379(6):547-56.

21.	Boughton CK, Bally L, Martignoni F, et al. Fully closed-loop insulin delivery in inpatients receiving nutritional support: a two-centre, open-label, randomised controlled trial. The Lancet Diabetes & Endocrinology. 2019;7(5):368-377.

Figure Legends

Figure 1 Panel a: A closed-loop system comprising a glucose sensor (black rectangle on the left-hand side of the abdomen), an insulin pump (device in the pocket), and a mobile-sized device containing the control algorithm (in user’s hand). Each component communicates with the other wirelessly (adapted from (22)). Panel b: The closed-loop system mimics the physiological feedback normally provided by the pancreatic beta-cell

Figure 2. Hybrid closed-loop system comprising fourth generation Enlite 3 glucose sensor, MiniMed® 670G insulin pump, with an embedded proportional-integral-derivative algorithm with insulin feedback (Medtronic, Northridge, CA) (permission from Medtronic).

Figure 3 The percentage of time that the glucose level was within the range of 70-180 mg/dL (3.9-10.0 mmol/L), as measured by continuous glucose monitoring, according to the time of day over 6 months among patients who were assigned to receive treatment with either the Tandem Control-IQ closed-loop system (closed loop) or a sensor-augmented pump (control). Symbols denote the hourly median values, and the shaded regions are defined by the 25th and 75th percentiles (reproduced with permission from (7)).