Role of PROPELLER-DWI of the prostate in reducing distortion and artefact from total hip replacement metalwork

# Marcin Czarniecki1,4, Iztok Caglic1, James Grist1, Andrew B. Gill1, Kamil Lorenc2, Rhys A. Slough1, Andrew N. Priest1, Tristan Barrett1,3.


1) Department of Radiology, Addenbrooke’s Hospital and University of Cambridge, Cambridge, UK

2) Nalecz Institute of Biocybernetics and Biomedical Engineering Polish Academy of Sciences, Warsaw, Poland
3) CamPARI Clinic, Addenbrooke’s Hospital and University of Cambridge, Cambridge, UK

4) Department of Radiology, Masovian Brodno Hospital, Warsaw, Poland


Corresponding author: 
Marcin Czarniecki (email: marcin.czarniecki@gmail.com, phone: +48223265531)
Address: Masovian Brodno Hospital

Department of Radiology

Ul. Kondratowicza 8

03-242 Warszawa

Abstract

Objective:

To compare image quality, artefact, and distortion in standard echo-planar imaging (EPI) with periodically rotated overlapping parallel lines with enhanced reconstruction (PROPELLER) for prostate magnetic resonance imaging (MRI) diffusion-weighted imaging (DWI) in patients with previous total hip replacement (THR).

Methods:

21 male subjects with a clinical suspicion for, or known prostate cancer and previous THR were scanned at 1.5 tesla using a phased-array body coil. DWI was obtained using single-shot EPI and PROPELLER techniques using fat saturation (PROPELLER-DWI-FS), and without (PROPELLER-DWI-NFS). Image quality (the overall impression of diagnostic quality) was compared to T2-weighted (T2WI) imaging using a 5-point Likert scale, with diffusion sequences additionally scored for artefact and distortion according to a 4-point scale, with artefact defined as the amount of prostate affected and distortion as the degree of warping of the organ. The T2W and DW image volumes were compared to produce quantitative distortion maps. A two-sample Wilcoxon test compared the qualitative scores, with inter-reader variability calculated using Cohen’s kappa.

Results:

21 patients were included in the study, with an average age of 70.4 years and PSA 9.2ng/mL. Hip metalwork was present bilaterally in 3 patients, left-sided in 9, and right-sided in 9. 
PROPELLER-DWI-FS significantly improved image quality (p<0.01) and reduced distortion (p<0.01) when compared to standard EP-DWI. Artefact was not shown to be significantly improved. The last 5 patients in the study were additionally imaged with PROPELLER-DWI-NFS, which resulted in a significant reduction in artefact compared to EP-DWI (p<0.05). Quantitative distortion was significantly lower compared to EP-DWI for both PROPELLER with fat saturation (p<0.01) and without fat saturation (p<0.01).

Conclusion:

PROPELLER-DWI demonstrates better image quality and decreases both artefact and distortion compared to conventional echo planar sequences in patients with hip metalwork.

Introduction

Multiparametric MRI with the inclusion of functional sequences has become a validated method for prostate lesion detection [1]. In particular, DWI has been shown to have an inverse correlation to tumour Gleason grade with the ability to discriminate between low-, intermediate, and high-risk prostate cancer [2–4]. Current reporting guidelines recommend DWI to be the key sequence for assessment of the peripheral zone, in which 75-80% of clinically significant prostate tumours are located [5,6].

DWI commonly utilizes magnetisation preparation gradients to sensitise the NMR signal to water motion, followed by a rapid echo planar imaging (EPI) readout [7]. However, the EPI method has limitations that can lead to numerous artefacts. Namely, hip metalwork may distort the B0, which in turn, leads to signal drop out, due to the subsequent rapid T2* decay induced in local tissues surrounding the implant as well as producing geometric distortion. Furthermore, the B0 shift may induce a frequency shift in the water species, leading to nulling of signals by commonly used spectral-spatial excitation pulses with a narrow chemical shift bandwidth, usually by 2ppm or less.. In combination, these factors can lead to significant image distortion, which hinder clinical interpretation.

Periodically rotated overlapping parallel lines with enhanced reconstruction (PROPELLER) combines a fast-spin echo (FSE) technique with radial sampling of k-space, producing a blade-like acquisition [13]. Susceptibility effects primarily affect T2* signal decay, by inducing local distortions in the static magnetic field. Therefore, with the use of a spin-echo pulse prior to each read out to refocus T2 whilst removing distorted T2* signal components from the NMR signal, PROPELLER sequences may decrease susceptibility induced subsequent image distortion [7, 8]. Furthermore, a radial acquisition has the advantage of oversampling the central portion of k-space, allowing bulk motion artefacts to be corrected in the reconstruction process. This allows it to be used for DWI without the distortions commonly associated with EPI readouts, thus its application may be helpful in decreasing distortion from metallic implants. PROPELLER-DWI has been shown to be effective in decreasing motion artefact in the brain, head and neck, musculoskeletal, and body MRI [8–13]. In prostate MRI, issues with motion artefact occur due to peristalsis or rectal spasm, but are less likely to deem a study uninterpretable when compared to metal artefacts which arise from metallic hip implants. The use of PROPELLER in diminishing metal artefact is not as well established, and was the primary purpose of this study.

Total hip replacement is a common orthopaedic operation in the developed world due to an ageing global population, with the rate predicted to double by 2035 [14]. Multi-parametric MRI (mpMRI) is increasingly being used to investigate suspected prostate cancer in this age group and therefore brings an increasing necessity for effective imaging techniques to help limit image distortion and artefact. The application of PROPELLER imaging in prostate imaging has been explored previously for T2W image improvement, and in a limited setting relating to DWI [9, 10], but there are no prior reports of DWI to reduce metalwork artefact. Therefore, in this study, we aimed to compare PROPELLER-DWI to standard echo planar (EP-DWI) sequences for the ability to improve image quality and reduce artefact and distortion in patients with hip metalwork.

Materials and methods

Study population

This single-institution retrospective study was approved by the institutional review board, with the need for informed consent for data analysis waived. 21 consecutive patients with previous THR who underwent 1.5T prostate MRI including EP-DWI and PROPELLER-DWI, were included and anonymised.

Image acquisition

Sequences were obtained on a 1.5T MR450 scanner (GE Healthcare, Waukesha, WI, USA) with a 16-channel cardiac phased array coil. All examinations included an axial Fast Spin-Echo T2WI sequence; repetition time (TR)/ echo time (TE) 5048/73ms, Field of View (FOV) 240cm, slice thickness/gap 4.5/0 mm, number of excitations (NEX) 2, matrix size 352x256, axial echo-planar diffusion-weighted imaging; TR/TE 3400/62ms, FOV 240mm, slice thickness/gap 4/0.5 mm, b-value 0, 1000 and axial PROPELLER-DWI with fat-saturation (PROPELLER-DWI-FS); TR/TE 5507/73, FOV 280mm, b-value 0, 800. All sequences were matched to T2WI in terms of slice thickness and orientation as per PI-RADS guidelines. 5 patients had an additional acquisition of PROPELLER-DWI without fat saturation (PROPELLER-DWI-NFS), TR/TE 5278/73.5, FOV 280mm (Table 1). PROPELLER was optimised by selecting a bandwidth to minimise TE, with all using 2D acquisition. Otherwise minimal changes were possible due to manufacturers restrictions. Slightly different FOV and b-value settings were used between EP-DWI and PROPELLER-DWI to optimise signal recoup. Standard axial T1WI, multi-planar T2WI and DCE imaging were also obtained, but were not analysed as part of this study.

### Image analysis

The images were independently reviewed and scored by two board-certified radiologists (X.X. and X. X.), with 7 and 2 years of reporting prostate MRI, respectively. T2WI, EP-DWI, PROPELLER-DWI-FS, and PROPELLER-DWI-NFS were reviewed in separate sessions, at least two weeks apart, in a random order with the readers being blinded to clinical data and patient history.

The radiologists scored T2WI, EP-DWI and PROPELLER-DWI-FS and PROPELLER-DWI-NFS for quality according to a 5-point scale: 1, poor; 2, suboptimal; 3, adequate for diagnosis; 4, above average; 5, excellent. The perceived image quality was used to describe the diagnostic quality of the obtained images, scores 1-2 were considered non-diagnostic, and scores of ≥3 considered adequate for diagnosis. All diffusion sequences were then scored for artefact (e.g. ghosting artefact, blurring, wrap or ring artefact) using a 4-point Likert scale [1]: 1, nil; 2, artefact visible in <½ of prostate transversely (right to left) or longitudinally (superior-inferiorly); 3, artefact in >½ of prostate transversely OR longitudinally; 4, artefact in >½ of prostate transversely AND longitudinally. A 4-point Likert scale was also used to score distortion: 1, mild mismatch covering <½ of the prostate with no warping; 2, involving <½ prostate transversely or longitudinally; 3, warping in >½ prostate transversely OR longitudinally; 4, warping in >½ prostate transversely AND longitudinally.

T2WI was used as the gold standard reference to assess distortion.

Outlining for quantitative analysis

Distortion in DWI images was evaluated relative to the T2WI ‘gold-standard’ image set using custom software written in MATLAB (Natick, MA) (version 2017a). Whole gland prostate outlining was performed by a single uroradiologist (X.X.). Outlines were first drawn on T2WI axial images, in a random patient order and blinded to the clinical information. Outlines were subsequently drawn on the relevant ADC maps in separate sittings for each of the DWI sequences, using the relevant T2WI axial images to aid anatomical localization, and to ensure equivalent slices were outlined. In all cases outlines could be determined on the relevant ADC maps, using T2WI as reference. A computational algorithm was then used to produce a measure of DW image distortion across the prostate, as previously described [17]. Briefly, a measure of the total mean distortion over the whole organ was formed by calculating the root-mean-square (r.m.s.) value of the magnitude of a distortion vector summed over all slices and angles.

Statistical analysis

Statistical analysis was performed in R language and environment for statistical computing (v3.3.2; R Foundation, Vienna, Austria). Inter-reader agreement was computed as a Cohen’s weighted kappa (κ). Agreement was interpreted as: Poor (κ<0.20); Fair (κ=0.21:0.40); Moderate (κ=0.41:0.60); Good (κ=0.61:0.80); and very good (κ=0.81:1.00) [18]. Image quality for T2WI, EP-DWI and PROPELLER-DWI were compared with a two-sample Wilcoxon test. Also, image artefact and distortion for EP-DWI and PROPELLER-DWI were compared with the same statistical test. All these tests were accounted for multiple comparisons with Holm-Bonferroni correction. Subsequently EP-DWI, PROPELLER-DWI-FS and PROPELLER-DWI-NFS were compared with a Friedman test using post-hoc analysis [20]. Root-mean-square value of the magnitude of distortion for EP-DWI, PROPELLER-DWI-FS and PROPELLER-DWI-NFS were compared using a two-sample Wilcoxon test.

### Results

21 patients were included in the study, with an average age of 70.4 years (range: 53-81) and PSA 9.2ng/mL (range: 2.3–25.1). All MRI studies included EPI and PROPELLER-DWI-FS, with 5 patients having additional PROPELLER-DWI-NFS. The indications for MRI were either primary detection (n=6), staging (n=3) or active surveillance (n=12). The final biopsy histopathological results were reported as: no cancer (n =4), Gleason 3+3 (n =8), 3+4 (n =5), 4+3 (n=3), 4+5 (n=1). Hip metalwork was present bilaterally in 3 patients, on the left side in 9, and right side in 9. For the 5 patients with additional PROPELLER-DWI-NFS imaging metalwork was present bilaterally in 1, on the left side in 3, and right side in 1.

T2-weighted Imaging

There was no statistically significant difference in scoring for T2WI quality (score  SD), reader 1: 4.7/5 ( 0.6) and reader 2: 4.3 ( 0.9), p = 0.18. EP-DWI image quality was scored as significantly lower than T2WI by both reader 1 1.6 ( 0.8), p<0.01 and reader 2 1.6 ( 0.8), p<0.01. Both readers also scored PROPELLER-DWI-FS on image quality as significantly lower than T2WI, reader 1: 2.5 ( 1.2), p<0.01, reader 2: 2.6 ( 0.9), p<0.01.

Propeller-DWI-FS vs. EP-DWI

PROPELLER-DWI-FS(n=21) exhibited significantly better image quality than EP-DWI for reader 1 (score  SD): 2.5 ( 1.2) vs. 1.6 ( 0.8), p < 0.05 and reader 2: 2.6 ( 0.9) vs. 1.6 ( 0.8), p < 0.01. Artefact was slightly decreased for PROPELLER compared to EP-DWI for reader 1: 2.9 ( 0.9) vs. 3.2 ( 0.9) and reader 2: 2.9 ( 1.0) vs. 3.1 ( 1.1), however this difference was not statistically significant (p=0.36 and p=1.0, respectively). Distortion was shown to decrease significantly for PROPELLER-DWI-FS by both reader 1: 2.0 ( 0.8) vs 3.8 ( 0.6), p<0.01 and reader 2: 1.4 ( 0.8) vs. 3.9 ( 0.4), p<0.01. The results are summarised in Table 2.

Propeller-DWI-NFS vs. EP-DWI

PROPELLER-DWI-NFS (n=5) was found to have significantly better image quality (2.9 ± 0.7 vs. 1.7 ± 0.9, p<0.05), less artefact (1.5 ± 0.5 vs. 2.9 ± 1.0, p<0.05), and distortion (1.0 ± 0.0 vs. 3.6 ± 0.7, p<0.05) in comparison to EP-DWI. The results are summarised in Table 2.

Propeller-DWI-NFS vs. Propeller-DWI-FS

PROPELLER-DWI-NFS, compared to PROPELLER-DWI-FS, showed a non-significant improvement in image quality (2.9 ± 0.7 vs. 2.5 ± 1.0 p=0.23). Distortion scores showed no significant difference (1.0 ± 0.0 vs. 1.6 ± 0.8, p=0.45). Image artefact was shown to significantly decreased (1.5 ± 0.5 vs. 2.9 ± 0.9, p<0.05), Table 3.

Inter-reader variability

There was moderate agreement (κ= 0.53) between readers for T2WI quality and very good agreement (κ= 0.94) for T2WI artefact assessment. EP-DWI demonstrated very good agreement (κ= 0.82) for scoring artefact, with good agreement (κ= 0.73) agreement for image quality. PROPELLER-DWI-FS had good agreement for both image quality (κ= 0.75) and artefact (κ= 0.78). There was poor (κ= 0.20) and fair (κ= 0.30) agreement between the two readers for assessing distortion in EP-DWI and PROPELLER-DWI-FS, respectively. PROPELLER-DWI-NFS (n=5) had full agreement (κ=1.00) for the two readers assessing distortion, good agreement in artefact scoring (κ= 0.62), and poor agreement on quality (κ= -0.04). The reader scoring and inter-reader variability parameters are summarised in Tables 2 and 3.

### Quantitative analysis of artefact and distortion

The root mean square resultant distortion for all (n=21) EP-DWI cases was 8.94 mm (range: 4.61–17.14 mm), PROPELLER-DWI-FS 3.70 mm (range: 1.95–6.24 mm) and 3.24 mm (range: 2.37–4.31 mm) for PROPELLER-NFS. The difference between root mean square resultant distortion between EP-DWI vs PROPELLER-DWI-FS and EP-DWI vs PROPELLER-DWI-NFS were found to be = significantly different (p<0.01 in both cases). The root mean square resultant distortion for PROPELLER-DWI-FS vs PROPELLER-DWI-NFS was not found to be significantly different (p = 1.0).  The individual results for all cases are listed in Table 4.

Discussion

PROPELLER imaging allows for FSE radial acquisition of k-space, leading to a decrease in artefact and distortion, and is theoretically advantageous for limiting susceptibility artefact induced by metalwork. Our study has shown that PROPELLER significantly improved imaging quality of DWI and reduced distortion from metallic implants compared to standard EP-DWI in patients with hip metalwork, with a significant reduction in artefact achieved by acquiring PROPELLER-DWI-NFS.
 	Previous studies utilizing PROPELLER imaging in the pelvis are limited, but its use has been shown to be effective in diminishing motion correction in anatomical regions outside of the prostate, including the male and female pelvis [16,20]. A similar feasibility study performed by Deng et al in the abdomen was able to show that PROPELLER-DWI acquisition improved image quality, decreased artefact and reduced geometric distortion, but study was based on a small patient group, and partly on phantoms [10]. In prostate MRI, PROPELLER has been verified as a viable technique for T2 weighted imaging in patients, but not for DWI, which is the key sequence for tumour detection in the peripheral zone [15]. A previous study on five healthy subjects by Roberts et al highlighted the potential benefit of PROPELLER DWI for prostate conspicuity by reducing artefacts associated with magnetic field inhomogeneities [16].  None of these previous studies assessed artefact reduction from metalwork, which was our primary purpose. Our study further confirmed previous results using a cohort of THR patients which was not studied before, with image distortion significantly improving, and a trend towards producing better quality images with less artefact. These qualitative results were also confirmed by quantitative assessment of the degree of distortion and supports earlier exploratory work in this area [17]. These findings can be attributed to PROPELLER’s unique radial k-space acquisition sequencing, which diminishes artefact in the phase-encoding direction. It should be noted than DW-PROPELLER is a proprietary method for implementing a FSE-based diffusion-weighted sequence, the fact that PROPELLER is FSE-based sequence, as opposed to EPI-based, leads to reduced metal susceptibility artefact. Other manufacturers offer similar, Cartesian diffusion-weighted TSE (DW-TSE or DW-HASTE) and the findings would be expected to pertain to Cartesian as well as radial acquisitions.

In the subgroup of patients undergoing PROPELLER-DWI-NFS image acquisition, there was a statistically significant improvement in scores for all three categories of image quality, artefact, and distortion. Although the sample size is small, this suggests that acquisition of PROPELLER without fat saturation may further improve results. It is hypothesised that the added fat-saturation pulse causes an additional signal drop-out in regions of off-resonance that may be produced in areas adjacent to the hip implant.

The inter-reader variability remained good or very good in categories of assessing image quality and artefact, with the one exception being the interpretation of image distortion. This may be explained by a difficulty in distinguishing artefact from distortion, as well as for image quality; the parameters are not necessarily mutually exclusive, potentially making the distinction more subjective and experience dependent. For this reason, the use of a quantifiable method of measuring distortion as outlined in this article and by Gill et al therefore may be useful in future studies to measure image distortion and to verify quality improvement. Using this method, we clearly demonstrated an improvement in distortion when comparing EP-DWI to PROPELLER-DWI images.

Our study has a number of limitations including its retrospective nature and limited sample size, in particular for the comparison of PROPELLER images with and without fat saturation which was performed in only 5 cases. The patient population was heterogeneous in terms of THR side, however the majority (18/21) were unilateral in nature. Although all patients had THR, the model and type of metalwork was also not correlated, and the results are not necessarily applicable to patients with other pelvic metalwork such as hip resurfacing, dynamic hip screws or fixation plates. Additionally, the different appearance of the DWI sequences meant blinding of these were not possible, which may have influenced scoring. Separate reading sessions avoided any bias relating to scoring the DWI sequences in direct opposition. We also did not evaluate or compare quantitative ADC values between the different sequences. It is likely however that these would give differing results due to the image distortion.

Although the routine use of PROPELLER is not currently recommended for general prostate DWI, its application in the subpopulation of patients following total hip replacement may be warranted. Pelvic imaging in these patients is often non-diagnostic in current clinical practice using EP-DWI, whilst the use of PROPELLER-DWI has been shown to increase image quality and decrease artefact and distortion. As noted in previous studies by Froehlich et al and Rosenkrantz et al this may be of importance for tumour diagnosis, as previous studies have shown a significant decrease in contrast [20]. In this study, we did not assess lesion conspicuity but this could be examined in future work.

Conclusion

Image acquisition using the PROPELLER technique helps provide images of improved quality, with decreased artefact and distortion when compared to the standard echo-planar DWI acquisition in patients following total hip replacement.

Acknowledgements

The authors acknowledge research support from National Institute of Health Research Cambridge Biomedical Research Centre, Cancer Research UK (CRUK), CRUK Cambridge Centre, CRUK and the Engineering and Physical Sciences Research Council (EPSRC) Imaging Centre in Cambridge and Manchester, and the Cambridge Experimental Cancer Medicine Centre (ECMC).

References

[1]	H.U. Ahmed, A. El-Shater Bosaily, L.C. Brown, R. Gabe, R. Kaplan, M.K. Parmar, Y. Collaco-Moraes, K. Ward, R.G. Hindley, A. Freeman, A.P. Kirkham, R. Oldroyd, C. Parker, M. Emberton, Diagnostic accuracy of multi-parametric MRI and TRUS biopsy in prostate cancer (PROMIS): a paired validating confirmatory study, Lancet. (2017). doi:10.1016/S0140-6736(16)32401-1.

[2]	T. Hambrock, D.M. Somford, H.J. Huisman, I.M. van Oort, J.A. Witjes, C. a Hulsbergen-van de Kaa, T. Scheenen, J.O. Barentsz, Relationship between apparent diffusion coefficients at 3.0-T MR imaging and Gleason grade in peripheral zone prostate cancer., Radiology. 259 (2011) 453–461. doi:10.1148/radiol.11091409.

[3]	O.F. Donati, Y. Mazaheri, A. Afaq, H.A. Vargas, J. Zheng, C.S. Moskowitz, H. Hricak, O. Akin, Prostate cancer aggressiveness: assessment with whole-lesion histogram analysis of the apparent diffusion coefficient., Radiology. 271 (2014) 143–52. doi:10.1148/radiol.13130973.

[4]	T. Barrett, A.N. Priest, E.M. Lawrence, D.A. Goldman, A.Y. Warren, V.J. Gnanapragasam, E. Sala, F.A. Gallagher, Ratio of tumor to normal prostate tissue apparent diffusion coefficient as a method for quantifying DWI of the prostate, Am. J. Roentgenol. 205 (2015) W585–W593. doi:10.2214/AJR.15.14338.

[5]	J.J. Lee, I.-C. Thomas, R. Nolley, M. Ferrari, J.D. Brooks, J.T. Leppert, Biologic differences between peripheral and transition zone prostate cancer., Prostate. 75 (2015) 183–90. doi:10.1002/pros.22903.

[6]	T. Barrett, B. Turkbey, P.L. Choyke, PI-RADS version 2: what you need to know, Clin. Radiol. (2015). doi:10.1016/j.crad.2015.06.093.

[7]	D. Le Bihan, C. Poupon, A. Amadon, F. Lethimonnier, Artifacts and pitfalls in diffusion MRI, J. Magn. Reson. Imaging. (2006). doi:10.1002/jmri.20683.

[8]	K.P. Forbes, J.G. Pipe, C.R. Bird, J.E. Heiserman, PROPELLER MRI: clinical testing of a novel technique for quantification and compensation of head motion., J. Magn. Reson. Imaging. 14 (2001) 215–222. doi:10.1002/jmri.1176.

[9]	C.-J. Juan, H.-C. Chang, C.-J. Hsueh, H.-S. Liu, Y.-C. Huang, H. Chung, C.-Y. Chen, H.-W. Kao, G.-S. Huang, Salivary Glands : Echo-Planar versus PROPELLER Diffusion- weighted MR Imaging for Assessment of ADCs, Radiology. 253 (2009) 144. doi:10.1148/radiol.2531082228/-/DC1.

[10]	J. Deng, F.H. Miller, R. Salem, R. Omary, A.C. Larson, Multishot diffusion-weighted PROPELLER magnetic resonance imaging of the abdomen., Invest. Radiol. 41 (2006) 769–775. doi:10.1097/01.rli.0000236808.84746.95.

[11]	B.F. Lane, F.Q. Vandermeer, R.C. Oz, E.W. Irwin, A.B. McMillan, J.J. Wong-You-Cheong, Comparison of Sagittal T2-Weighted BLADE and Fast Spin-Echo MRI of the Female Pelvis for Motion Artifact and Lesion Detection, Am. J. Roentgenol. 197 (2011) W307–W313. doi:10.2214/AJR.10.5918.

[12]	K. Nagatomo, H. Yabuuchi, Y. Yamasaki, H. Narita, S. Kumazawa, T. Kojima, N. Sakai, M. Masaki, H. Kimura, Efficacy of periodically rotated overlapping parallel lines with enhanced reconstruction (PROPELLER) for shoulder magnetic resonance (MR) imaging, Eur. J. Radiol. 85 (2016) 1735–1743. doi:10.1016/j.ejrad.2016.07.008.

[13]	J.G. Pipe, Motion correction with PROPELLER MRI: Application to head motion and free-breathing cardiac imaging, Magn. Reson. Med. 42 (1999) 963–969. doi: 10.1002/(SICI)1522-2594(199911)42:5<963::AID-MRM17>3.0.CO;2-L

[14]	National Joint Registry for England Wales and Northern Ireland, 13th Annual Report, 2016. http://www.njrcentre.org.uk/njrcentre/Portals/0/Documents/England/Reports/13th Annual Report/07950 NJR Annual Report 2016 ONLINE REPORT.pdf.

[15]	A.B. Rosenkrantz, G.L. Bennett, A. Doshi, F.M. Deng, J.S. Babb, S.S. Taneja, T2-weighted imaging of the prostate: Impact of the BLADE technique on image quality and tumor assessment, Abdom. Imaging. 40 (2014) 552–559. doi:10.1007/s00261-014-0225-7.

[16]	T.P. Roberts, M. Haider, Diffusion weighted imaging of the Prostate Gland in the Face of Magnetic Susceptibility Differences – Parallel EPI and PROPELLER FSE approaches, Proc. Intl. Soc. Mag. Reson. Med. 11 (2004) 946.

[17]	A.B. Gill, M. Czarniecki, F.A. Gallagher, T. Barrett, A method for mapping and quantifying whole organ diffusion-weighted image distortion in MR imaging of the prostate, Sci. Rep. 7 (2017) 12727. doi:10.1038/s41598-017-13097-6.

[18]	J. Cohen, Weighted kappa: nominal scale agreement with provision for scaled disagreement or partial credit., Psychol. Bull. 70 (1968) 213–220. doi:10.1037/h0026256.

[19]	T. Galili, Post hoc analysis for Friedman’s Test (R code), R-Statistics Blogs. (2010). https://www.r-statistics.com/2010/02/post-hoc-analysis-for-friedmans-test-r-code/ (accessed February 6, 2017).

[20]	J.M. Froehlich, T. Metens, B. Chilla, N. Hauser, M. Klarhoefer, R.A. Kubik-Huch, Should less motion sensitive T2-weighted BLADE TSE replace Cartesian TSE for female pelvic MRI?, Insights Imaging. 3 (2012) 611–618. doi:10.1007/s13244-012-0193-9.

Figure 1. Box plot of measured values as the root mean square (r.m.s) of resultant distortion (mm) between EP-DWI (DWI), PROPELLER-DWI-FS (PROPELLER DWI with fat saturation) and PROPELLER-DWI-NFS (PROPELLER DWI without fat saturation)

Figure 2. Example volume distortion maps for quantitative analysis of DWI.
74-year old patient with left THR, PSA 5.21ng/ml, on active surveillance for Gleason 3+3 disease. No lesion on MRI. A, C, E: ADC maps from different DWI acquisitions. B, D, F: Respective quantitative distortion maps. A: ADC map from EP-DWI shows significant warping and distortion, confirmed on quantitative distortion map (B). C: ADC map from PROPELLER-DWI-FS shows mild ring artefact, but little distortion of outline (D). E: ADC map from PROPELLER-DWI-NFS has little artefact or distortion (F). (Blue regions in distortion maps correspond to ‘recesses’ in the DWI volume compared to the T2W volume; red areas correspond to ‘bulges’. The vector tree shown in B, D & F shows the displacement of individual slice centroids with respect to the T2W image volume.)

Figure 3. Reduced distortion on PROPELLER DWI.

A, B: 74-year old patient, PSA 7.3 ng/ml. A: T2 weighted image shows a large 25x15 mm anterior midline TZ lesion (*). B: ADC map from EP-DWI shows restricted diffusion, but warping and distortion due to right THR makes direct matching challenging. C: ADC map from PROPELLER-DWI-FS shows no distortion and good correlation of restricted diffusion compared to T2 lesion; transperineal biopsy in this region demonstrated Gleason 3+4 disease.

Figure 4. Improved lesion assessment with PROPELLER-DWI-FS.

72-year-old patient, PSA 14.1 ng/ml. A: T2 weighted image shows a 10 mm area of focal low signal in the right mid/apex PZ (arrow). B: ADC map from EP-DWI in this region is not diagnostic due to warping from right THR. C: ADC map from PROPELLER-DWI-FS with no distortion shows focal restriction in this lesion (arrow) consistent with a high probability lesion; subsequent targeted biopsy confirmed Gleason 3+4 disease in 30% of cores.

Figure 5. Reduced artefact at non-fat saturated PROPELLER DWI

72-year-old patient on active surveillance for Gleason 3+3 disease. A: T2 image at the base shows artefact form left THR not extending to prostate. B: ADC map from EP-DWI shows artefact involving more than half the prostate in the transverse and longitudinal planes and significant distortion. C: ADC map from PROPELLER-DWI-FS shows ring artefact over the prostate but no distortion of outline. D: ADC map from PROPELLER-DWI-NFShas only mild artefact or distortion at the same level.

Table 1. Magnetic resonance imaging (MRI) protocols

| Sequence name |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Fast Spin-Echo T2-Weighted imaging | Echo Planar Diffusion-Weighted imaging | PROPELLER DWI | PROPELLER DWI (without fat saturation) |
| TR/TE (ms) | 5048/73 | 3400/62 | 5507/73 | 5278/73 |
| FOV (cm) | 240 | 240 | 280 | 280 |
| b-values | N/A | 0, 1000 | 0, 800 | 0, 800 |
| Fat suppression | No | No (spectral spatial excitation) | Yes | No |
| Echo-train length | 16 | Single shot | 16 | 16 |
| Number of excitations | 2 | 6 | 2 | 2.5 |
| Acquisition Matrix | 352x256 | 128x128 | 128x128 | 128x128 |
| Acquisition time (mean) | 5:59 | 2:58 | 6:46 | 5:46 |

Table 2. Reader scoring

|  | Reader 1 (mean ± SD) |  |  |  |
| --- | --- | --- | --- | --- |
|  | T2WI | EP-DWI | PROPELLER-DWI-FS | PROPELLER-DWI-NFS |
| All patients (n=21) |  |  |  |  |
| Quality | 4.7 ± 0.6 | 1.6 ± 0.8 | 2.5 ± 1.2† |  |
| Artefact | 1.4 ± 0.7 | 3.2 ± 0.9 | 2.9 ± 0.9† |  |
| Distortion |  | 3.8 ± 0.6 | 2.0 ± 0.8† |  |
| Propeller-DWI-FS vs. EP-DWI (n=5) |  |  |  |  |
| Quality | 5.0 ± 0.0 | 1.4 ± 0.9 | 2.8 ± 1.4 | 3.8 ± 0.9† |
| Artefact | 1.4 ± 0.6 | 3.2 ± 0.9 | 2.6 ± 0.9 | 1.4 ± 0.6†‡ |
| Distortion |  | 3.6 ± 0.9 | 1.6 ± 0.9 | 1.0 ± 0.0† |
|  | Reader 2 (mean ± SD) |  |  |  |
|  | T2WI | EP-DWI | PROPELLER-DWI-FS | PROPELLER-DWI-NFS |
| All patients (n=21) |  |  |  |  |
| Quality | 4.3 ± 0.9 | 1.6 ± 0.8 | 2.6 ± 0.9† |  |
| Artefact | 1.0 ± 0.0 | 3.1 ± 1.1 | 2.9 ± 1.0 |  |
| Distortion |  | 3.9 ± 0.4 | 1.4 ± 0.8† |  |
| Propeller-DWI-FS vs. EP-DWI (n=5) |  |  |  |  |
| Quality | 4.0 ± 0.0 | 1.4 ± 0.6 | 2.6 ± 0.6 | 3 ± 0.7† |
| Artefact | 1.0 ± 0.0 | 3.0 ± 1.0 | 2.8 ± 1.1 | 1.6 ± 0.6 ‡ |
| Distortion |  | 4.0 ± 1.0 | 1.0 ± 0.0 | 1.0 ± 0.0† |
| † Statistically significant when compared to EP-DWI using a two-sample Wilcoxon score test |  |  |  |  |
| ‡ Statistically significant when compared to PROPELLER-DWI-FS using a two-sample Wilcoxon score test Note: Statistical tests in the table were not corrected for multiple comparisons, due to the small population that was studied. |  |  |  |  |

Table 3. Inter reader variability (Cohen’s kappa)

| Scoring (kappa ± 95% CI) | EP-DWI n=21 | PROPELLER-DWI-FS n=21 | PROPELLER-DWI-NFS n=5 |
| --- | --- | --- | --- |
| Quality | 0.73 (0.40 - 1.1) | 0.75 (0.63 - 0.87) | -0.040 (-1.22 - 0.417) |
| Artefact | 0.82 (0.70 - 0.94) | 0.78 (0.62 - 0.94) | 0.62 (-0.007 - 1.2) |
| Distortion | 0.20 (-0.37 - 0.02) | 0.30 (-0.02 - 0.63) † | 1.0 (1.0 - 1.0) |
| † Interreader variability is statistically significant |  |  |  |

Table 4. Root mean square distortion (mm) of prostate attained using computational algorithm

| Case | EP-DWI | PROPELLER-DWI | PROPELLER-DWI-NFS |
| --- | --- | --- | --- |
| 1 | 6.79 | 3.01 | - |
| 2 | 12.42 | 4.49 | - |
| 3 | 8.62 | 4.02 | - |
| 4 | 4.22 | 2.78 | - |
| 5 | 4.64 | 1.84 | - |
| 6 | 14.21 | 2.85 | - |
| 7 | 16.49 | 3.58 | - |
| 8 | 4.61 | 1.95 | - |
| 9 | 4.62 | 4.10 | - |
| 10 | 4.88 | 4.13 | - |
| 11 | 13.68 | 6.24 | - |
| 12 | 4.21 | 3.71 | - |
| 13 | 11.36 | 2.91 | - |
| 14 | 17.14 | 5.55 | - |
| 15 | 7.25 | 5.65 | - |
| 16 | 6.42 | 3.59 | - |
| 17 | 4.92 | 4.22 | 3.49 |
| 18 | 10.31 | 3.38 | 2.37 |
| 19 | 12.79 | 2.92 | 4.31 |
| 20 | 8.99 | 3.92 | 2.91 |
| 21 | 9.16 | 2.92 | 3.12 |
| Group mean | 8.94 | 3.70 | 3.24 |