Improving the safety and tolerability of local anaesthetic out-patient transperineal prostate biopsies: a pilot study of the CAMbridge PROstate Biopsy (CAMPROBE) method

D Thurtle1,2, L Starling1,2,3, K Leonard1,2,3, T Stone4, VJ Gnanapragasam1,2,3

1 – Academic Urology Group, University of Cambridge, UK,

2 – Department of Urology, Addenbrooke’s Hospital, Cambridge, Cambridge, UK

3 – Cambridge Urology Translational Research and Clinical Trials, University of Cambridge, UK

4 – Medical Physics and Clinical Engineering, Addenbrooke’s Hospital, Cambridge, UK

Corresponding author:

Vincent J Gnanapragasam: vjg29@cam.ac.uk

Academic Urology Group

Box Number 279
Addenbrooke’s Hospital

Hills Road
Cambridge CB2 0QQ

Tel: 01223763365
Fax: 01223769007

Abstract

Objectives
To pilot the use of a bespoke device (CAMPROBE) to enable routine out-patient free-hand local anaesthetic (LA) transperineal prostate biopsies.

Patients and methods

The CAMPROBE prototype was designed and built in our institution. Men on active surveillance due prostate re-sampling were invited to have a CAMPROBE biopsy as an alternative to repeat transrectal biopsies (TRUSBx) as part of an approved trial (NCT02375035). Biopsies were performed using LA infiltration only, without sedation or additional analgesia. Patient reported outcomes were recorded at Day 0 and 7 using validated questionnaires and visual analogue scales (VAS). Complications were recorded prospectively.

Results
30 men underwent biopsies with a median of 11 cores taken per procedure (IQR 10-12). There were no infections, sepsis or retention episodes. Haematuria and haematospermia occurred in 67% and 62% which is similar to rates reported for TRUSBx. Mean VAS for pain (0-10 scale) were less than 3 for every part of the procedure. All 30 men described the procedure as tolerable under LA. 26/30 (86.7%) men expressed a preference for a CAMPROBE procedure over TRUSBx and a further 3 (10.0%) would have either.

Conclusions
In this small pilot study the CAMPROBE device and method appears to be a safe, simple and well-tolerated out-patient transperineal alternative to TRUSBx. A major new NIHR grant will allow its further development from a protoype to a single use low-cost disposable device ready for multi-centre testing.

Key words

Prostate cancer
Transperineal biopsy
Diagnostic intervention
Patient safety
Service delivery

Level of evidence
1b – Individual cohort study

Introduction
Prostate cancer (PCa) is the commonest male cancer, representing 25% of all male neoplasms in the UK (1). The mainstay of histological diagnosis remains a transrectal ultrasound guided prostate biopsies (TRUSBx).  Over 1 million TRUSBx are performed annually in the US and Europe (2). However, these are associated with unacceptably high risks of infection, with fevers or chills reported in over 10% and severe sepsis in 1-2% (3, 4). Costs associated with UK hospitalisation episodes for biopsy-related sepsis alone are estimated to be £7-11million annually (5). These costs and risks will be amplified by increasing referral trends, projected to rise two-thirds by 2030, and by worrying rises in antibiotic resistance of up to 4-fold (6, 7). Additionally, the transrectal approach has recognised drawbacks in undersampling the anterior and apical aspects of the gland (8). It is clear the TRUSbx approach is untenable as a long term option as a diagnostic procedure.

Transperineal biopsy (TPBx) has significantly lower infective risks as it traverses the sterilised perineum and in fact predates TRUSBx (9, 10). Techniques using a perineal brachytherapy template grid under general anesthesia (GA), or a free-hand ‘fan technique’ under local anaesthesia (LA) are both well established, although significant regional variations in practice exist (11). The use of general anaesthesia and specialised equipment for grid based TPBx limit its utility as a routine replacement for TRUSBx. Attempts at doing grid-base biopsies without GA have reported use of 50-60ml of local anaesthesia, additional analgesia and/or significant patient pre-preparation which is not compatible with routine out-patient or office based practice(12, 13).

The free-hand LA TPBx technique is appealing as it requires only 2 perineal punctures and no specialist equipment. However, the procedure is not standardised, nor commonly performed in UK centres. We sought a way to enable wider uptake of LA TPBx as a viable cost-effective and routine replacement for TRUSBx by developing the The CAMbridge PROstate Biopsy (CAMPROBE). The CAMPROBE (patent pending) is a cannulated TP access system based on the co-axial concept but bespoke to the context of prostate biopsies (14). The integrated device allows for synchronous device insertion and LA infiltration under ultrasound guidance, negating the need for separate punctures, nerve blocks or sedation. Once in position, standard 18G core-needle biopsies can be taken through the retained cannula. No specialized equipment except a couch suitable for lithotomy and a linear array ultrasound probe are required. Here we report clinical and patient reported outcomes form our pilot trial of the prototype device.

Patients and methods

CAMPROBE design and manufacture

The CAMPROBE was designed towards the aim of producing a simple device to facilitate TP biopsies under LA, with procedures performed by a single operator with 1 assistant. The device (figure 1) features an integrated needle to which a standard syringe can be attached for LA delivery. This is sheathed within a coaxial blunt ended cannula long enough to penetrate from perineal skin to prostate capsule. The proximal end of the sheath was designed as a taper to permit easy introduction of a standard biopsy needle. The device was designed and manufactured in collaboration with the Clinical Engineering Innovation team (Cambridge University Hospitals NHSFT), the design and manufacture met the requirements of the medical devices directive for in-house manufactured devices, this was assured within the departmental ISO 13485 certified design processes. The design process followed an iterative healthcare design process developed by Cambridge University, Engineering Design Centre (https://www-edc.eng.cam.ac.uk/research/healthcaredesign/).

Clinical cohort

Following the innovation phase, we commenced the prospective development phase (IDEAL Stage 2a) by piloting the prototype CAMPROBE device in an initial group of patients (15). Following trial approval and registration (NCT02375035) and ethical approval (REC reference: 14/EE/1172) men aged 18-85, who were already awaiting a prostate biopsy and had undergone previous transrectal biopsies, were recruited from a prostate cancer monitoring (active surveillance) programme. As this pilot trial was designed to primarily test safety and tolerability, we excluded men with suspicious mpMRI lesions who needed image guided biopsies.

Procedure

All procedures were undertaken in a clinic out-patient room.  A single surgeon (VJG) performed the biopsies with the assistance of 1 nurse or health care assistant. All men received a pre-biopsy dose of ciprofloxacin 500mg and had self-administered a glycerol suppository a few hours before the procedure. The patient was positioned in a modified lithotomy position on a perineal suturing couch with legs on holders (figure 2). The perineal area was prepared with Betadine or Chlorhexidine wash. Two points were marked 1.5cm above the anal verge and 1.5cm on either side of the midline (figure 3). At each marked site 0.5ml of 1% lignocaine was injected to numb the skin. The CAMPROBE was assembled and attached to a 5 or 10ml syringe filled with 1% lignocaine (figure 4a). A linear array US transducer was placed in the rectum and the prostate and perineal tissue visualised. Starting on the right, the CAMPROBE was inserted under vision with LA delivered to sub-cutaneous tissue and deep perineal muscle synchronously as the device was advanced.  As the device approached the prostate capsule about 2-3mls of LA was delivered to the peri-prostatic space. Once in position the syringe and integrated needle was removed leaving the co-axial access sheath in place (figure 4b). Through this sheath, standard 18G biopsies were taken (using a BARD® biopsy device with a 22cm needle) from the posterior and anterior prostate and mid and peripheral zones (a maximum of 6 on each side) using a fan approach (figures 5a and 5b). If there was any further discomfort, the integrated needle was replaced and further LA injected without the need for a further skin puncture. In total between 5-10mls of lignocaine/side was required depending on the patient’s tolerability. The US probe was rotated and angled to visualise the needle direction and biopsy site for each acquisition. In this respect the thicker CAMPROBE cannula permitted improved visualisation of the biopsy needle direction. This was repeated on the contralateral side with a total procedure time of between 15-30 minutes. Once the CAMPROBE is removed, 2-3 minutes of perineal pressure is applied followed by light dressings to each puncture site. Post-biopsy, patients were managed in the standard fashion after a TRUSBx and given 3 days of post biopsy oral antibiotics.

Outcomes analysis

The primary objective was to assess patient reported outcomes and tolerability of the procedure. Secondarily we sought to assess the safety of the procedure by assessing incidence of infections or biopsy-related healthcare encounters. Evaluation was achieved using questionnaires adapted from the ProBE study, embedded within the ProtecT trial (3). All questionnaires underwent review and were approved by the assigned Research Ethics Committee.  Men self-reported pain, comfort, embarrassment and dignity on a 4-point Likert-type scale immediately after, and seven days following CAMPROBE biopsy. Discomfort from each stage of the procedure, and attitudes towards re-biopsy were assessed using a 0-10 visual analogue scale (VAS). The complete questionnaires are available in the online supplementary files.   Complications or adverse events were self-reported at day 7, with the severity of complications assessed on a four-point scale from ‘none’ to ‘major’. For analysis, adverse events were classified according to the National Cancer Institute’s toxicity criteria which ranges from 0 – where no adverse event is reported, to grade 4 – death as a result of the biopsy (16) . Data was collated and analysed anonymously using Stata 14 (Texas, USA). There was no requirement to impute missing data.

Results

Study cohort

34 men were recruited to the study. Of these, 30 underwent the CAMPROBE procedure. 3 men were unable to adopt the lithotomy position due to hip abduction problems and 1 man was unwell with an arrhythmia prior to biopsy. Median age and PSA was 71.5 years and 5.3ng/ml respectively. There were no device failures or adverse events during the procedure. All men had undergone at least one previous TRUSBx (Table 1). All procedures were completed within 30 minutes. In total 334 biopsy cores were taken (median 11, IQR 10-12). Pathological results were as expected in this low-risk cohort, with 13 (43.3%) men found to have cancer in 1 or more biopsy cores and 4 men (13.3%) pathologically upgraded as a result of the biopsy.

Patient reported outcomes

Across the cohort, mean VAS scores for discomfort, on a 0-10 scale, for the initial digital rectal examination, scanner probe insertion, LA injection and biopsy-taking were 2.04 (SD, 1.9), 2.97 (SD, 1.9), 2.67 (SD, 2.2) and 1.83 (SD, 1.2) respectively (table 2). No one reported the biopsies to be the most painful part of the procedure. No procedures were abandoned due to pain.  29/30 men (96.7%) scored the procedure overall to be 4/10 or less for pain, with all 30 men describing the procedure as tolerable under LA. 27/30 men (90.0%) rated the CAMPROBE biopsy superior or equivalent to a transrectal biopsy. 20 and 9 men reported they would have ‘no problem’ and only a ‘minor problem’ with undergoing a repeated CAMPROBE biopsy respectively (table 3). 26/30 (86.7%) men would prefer a CAMPROBE procedure over a transrectal-biopsy for their next biopsy and a further 3(10.0%) would have either.

Clinical complications

There were no peri-procedural complications with all men discharged after the procedure. No men required hospital re-admission in the week following their biopsy, including zero episodes of urinary tract infection or systemic sepsis. One man (3.3%) reported seeing his general practitioner as a result of pain from the biopsy, and was prescribed painkillers. 20 men (66.7%) reported visible haematuria by day 7 and 12 men reported haematospermia amongst a cohort of 19 who reported an ejaculation in the week following biopsy (63.2%). Two men (6.7%) reported blood in their stools which persisted for 4 and 5 days respectively. Median duration of haematuria, when present, was 3 days, no one required medical attention for bleeding and no one reported a bleeding complication as more than ‘a minor problem’. 10 men (33.3%) reported alterations in urinary function in the week following the biopsy, with increased frequency and increased nocturia both reported, however no one reported this to be more than ‘a moderate problem’ and no one sought medical attention. There were no episodes of urinary retention. Rates of fever, shivering or infection, haematochezia, and biopsy-related health encounters were all much lower in our cohort compared to the ProBE study (table 4) (3). Haematuria and haematospermia rates were very similar.

Discussion
These early promising results from this pilot study suggest proof-of-principle that the CAMPROBE device and method for LA free-hand TP-biopsy approach is feasible, well tolerated by patients, and preferable to transrectal-biopsy for the vast majority of men.  Amongst the ProtecT study TRUSBx ProBE cohort 213/1085 (19.6%) reported that having a repeat transrectal-biopsy would represent a moderate/major problem, when asked at 7 days(3). This compares to only 1/30 (3.3%) reporting a moderate problem with having another CAMPROBE procedure. We report a very low complication risk – importantly with zero episodes of infection or sepsis – and biopsy-related health encounters of any type (3.3%) were also considerably lower than the 11.7% reported following TRUSBx within the ProBE study (3). Both of which if confirmed in larger trials would have major patient-safety and cost-reduction implications.

Because the CAMPROBE is an integrated device, LA is delivered under vision which allows precise deployment of LA and hence reduced dosages. In addition the free-hand approach makes the probe position much more comfortable for men negating the need for anal sphincter relaxants or rectal LA gel. These facts make the CAMPROBE method much safer as an out-patient technique with minimal staffing (akin to the standard TRUSBx). The method exploits already existing equipment and consumables in any urology unit and will thus be easy to adopt. Future device costs are also projected to be low and on a par with current costs for disposable needle guides used for TRUSBx.

In 2015 a meta-analysis reported no difference in cancer detection between free-hand TP-biopsy and transrectal-biopsy (17). However, to our knowledge this technique has not been widely adopted in the UK or elsewhere (18).  Most studies have performed free hand biopsies by simply inserting the biopsy needle into the perineal skin directly. In our view this can be technically challenging due to the flexibility of the unsupported biopsy needle and the need for repeated skin passes. Certainly the co-axial approach has been shown to significantly reduce patient discomfort from free hand TP biopsies (14).  Our method also goes further in reducing the need for separate punctures to deliver LA by integrating the LA delivery needle.  The device also simplifies the procedure for the operating surgeon, with the device able to reach the prostate directly and an ergonomic design with a wide proximal funnel for simple single-handed biopsy-taking.

Limitations to this pilot study exist. As with any questionnaire-based study limitations exist with regards to recall bias, particularly when comparisons are made to previous transrectal-biopsy which may have occurred many months earlier. However, the questionnaire response rate of 100% is a particular strength of the study. Completed VAS scores from 30 men is significantly higher than comparable publications drawing similar conclusions (12). Secondly in this study we have focussed on men without targetable lesions on mpMRI as we did not want to potentially compromise histological detection. However, the technology could already be used for cognitive image-guided biopsies and no difference in clinically significant cancer detection has yet been demonstrated between cognitive and MRI-USS fusion techniques in either a comparative trial or systematic review of the relevant literature (19, 20). Indeed in a recent paper Yaxley et al reported no difference in either overall or clinically significant cancer detection rates in men having either image-guided in-bore MRI biopsies, transperineal grid based biopsies or cognitively targeted TRUSBx (21). In this context the CAMPROBE technique would also be preferable to transrectal-biopsy due to the ability to access and sample the whole prostate gland. Image-guided techniques for use with free-hand TP-biopsy are also likely to be clinically available in the very near future, with the technology rapidly developing and 2 recent small studies demonstrating feasibility alongside the free-hand technique and suggesting superiority to systematic biopsy alone in detecting significant cancer (22, 23). Further to this we found that after device insertion, there was no association with increasing numbers of biopsies taken and patient discomfort. Thus we foresee no issues with obtaining both targeted biopsies and systematic sectoral biopsies using our method. In contrast, recent reports of grid based TP approaches under LA for cognitive image guided biopsies have only performed targeted biopsies presumably to try and limit patient discomfort (12).

The focus of this pilot study was to assess safety and tolerability rather than tumour detection. However, comparisons with recently published cohorts performing re-biopsy in a similar context provide reassurance that our method does not compromise tumour detection.  For example, repeat 12-core sectoral systematic biopsy detected PCa in 56% of men on AS with a negative mpMRI in a recent study from Johns Hopkins, Baltimore(24). In terms of detecting pathological tumour-upgrade, 13.3% of men were upgraded on CAMPROBE systematic biopsy. This is very similar to the 11% reported amongst those with no MRI lesion in a similar study from Memorial Sloan Kettering, New York(25). Tumour detection is clearly an important outcome, and will be thoroughly assessed in future clinical testing.

As this was a pilot feasibility study we do interpret our results with caution as we have small numbers, a lack of randomisation and no direct comparison to other biopsy types – including with respect to cancer-detection. However these will be addressed in a future multi-centre, large-volume (IDEAL Stage 3) study. To facilitate this we have recently secured a major ‘Invention for Innovation’ Product Development Award from the National Institute for Health Research (NIHR), which recognises the potential beneficial impact of this technique as a replacement for TRUSBx. Prospective NIHR funding over the next few years will enable the commercialisation of the CAMPROBE and transition from our surgical steel prototype to a disposable, CE-marked mass produced single-use device. This will further reduce costs and increase accessibility to the device and technique, potentially accelerating the introduction into widespread clinical practice. The future scope for benefit is therefore potentially enormous and this innovation could have a major and critical impact for patients worldwide. Undoubtedly, there is a learning curve with the free hand LA TP method as it involves dual hand co-ordination which is different from the discipline of TRUSBx. However free hand TP biopsy is a well known technique with a number of publications, suggesting it is a skill that can be acquired relatively quickly by a competent clinician (22, 26). Of note, inability to tolerate the lithotomy position may be an issue in a small proportion of men with 3/34 (8.8%) in our cohort unable to tolerate the position, predominantly due to hip discomfort.

In summary TRUSBx for suspected prostate cancer exposes men to excessive risk of infections and sepsis and contributes to an increasing crisis in antibiotic resistance. The CAMPROBE is a simple, cheap but disruptive technology that appears to be a good out-patient LA alternative, using the safer transperineal route. The simplicity and standardisation of the procedure should allow widespread uptake, with significant safety and cost-benefits for health service delivery. Our current NIHR-funded work to develop a cheap, disposable version for the device will accelerate the journey to clinical practice.

List of abbreviations

PCa – Prostate cancer

TP – transperineal

LA – local anaesthetic

NCT – National Clinical Trial

CAMPROBE – Cambridge Prostate Biopsy Device

References

1.	Cancer Research UK  [Available from: http://www.cancerresearchuk.org/health-professional/cancer-statistics/incidence/all-cancers-combined.

2.	Loeb S, Vellekoop A, Ahmed HU, et al. Systematic review of complications of prostate biopsy. Eur Urol. 2013;64(6):876-92.

3.	Rosario DJ, Lane JA, Metcalfe C, et al. Short term outcomes of prostate biopsy in men tested for cancer by prostate specific antigen: prospective evaluation within ProtecT study. BMJ. 2012;344:d7894.

4.	Anastasiadis E, van der Meulen J, Emberton M. Hospital admissions after transrectal ultrasound-guided biopsy of the prostate in men diagnosed with prostate cancer: a database analysis in England. Int J Urol. 2015;22(2):181-6.

5.	Batura D, Gopal Rao G. The national burden of infections after prostate biopsy in England and Wales: a wake-up call for better prevention. J Antimicrob Chemother. 2013;68(2):247-9.

6.	Mistry M, Parkin DM, Ahmad AS, et al. Cancer incidence in the United Kingdom: projections to the year 2030. Br J Cancer. 2011;105(11):1795-803.

7.	Carignan A, Roussy JF, Lapointe V, et al. Increasing risk of infectious complications after transrectal ultrasound-guided prostate biopsies: time to reassess antimicrobial prophylaxis? Eur Urol. 2012;62(3):453-9.

8.	Han M, Chang D, Kim C, et al. Geometric evaluation of systematic transrectal ultrasound guided prostate biopsy. J Urol. 2012;188(6):2404-9.

9.	Barringer B. Carcinoma of the prostate. Surg Gynecol Obstet. 1922;34:168–76.

10.	Grummet JP, Weerakoon M, Huang S, et al. Sepsis and 'superbugs': should we favour the transperineal over the transrectal approach for prostate biopsy? BJU Int. 2014;114(3):384-8.

11.	Galfano A, Novara G, Iafrate M, et al. Prostate Biopsy: The Transperineal Approach. European Association of Urology and European Board of Urology. 2007(EAU - EBU update series 5):241–9.

12.	Bass EJ, Donaldson IA, Freeman A, et al. Magnetic resonance imaging targeted transperineal prostate biopsy: a local anaesthetic approach. Prostate Cancer Prostatic Dis. 2017.

13.	Smith JB, Popert R, Nuttall MC, et al. Transperineal sector prostate biopsies: a local anesthetic outpatient technique. Urology. 2014;83(6):1344-9.

14.	Novella G, Ficarra V, Galfano A, et al. Pain assessment after original transperineal prostate biopsy using a coaxial needle. Urology. 2003;62(4):689-92.

15.	McCulloch P, Altman DG, Campbell WB, et al. No surgical innovation without evaluation: the IDEAL recommendations. Lancet. 2009;374(9695):1105-12.

16.	Trotti A, Bentzen SM. The need for adverse effects reporting standards in oncology clinical trials. J Clin Oncol. 2004;22(1):19-22.

17.	Shen PF, Zhu YC, Wei WR, et al. The results of transperineal versus transrectal prostate biopsy: a systematic review and meta-analysis. Asian J Androl. 2012;14(2):310-5.

18.	Dundee PE, Grummet JP, Murphy DG. Transperineal prostate biopsy: template-guided or freehand? BJU Int. 2015;115(5):681-3.

19.	Wysock JS, Rosenkrantz AB, Huang WC, et al. A prospective, blinded comparison of magnetic resonance (MR) imaging-ultrasound fusion and visual estimation in the performance of MR-targeted prostate biopsy: the PROFUS trial. Eur Urol. 2014;66(2):343-51.

20.	Wegelin O, van Melick HH, Hooft L, et al. Comparing Three Different Techniques for Magnetic Resonance Imaging-targeted Prostate Biopsies: A Systematic Review of In-bore versus Magnetic Resonance Imaging-transrectal Ultrasound fusion versus Cognitive Registration. Is There a Preferred Technique? Eur Urol. 2017;71(4):517-31.

21.	Yaxley AJ, Yaxley JW, Thangasamy I,et al. Comparison between target MRI in-gantry and cognitive target transperineal or transrectal guided prostate biopsies for PIRADS 3-5 MRI lesions. BJU Int. 2017.

22.	Zhang Q, Wang W, Yang R, et al. Free-hand transperineal targeted prostate biopsy with real-time fusion imaging of multiparametric magnetic resonance imaging and transrectal ultrasound: single-center experience in China. Int Urol Nephrol. 2015;47(5):727-33.

23.	Lian H, Zhuang J, Wang W, et al. Assessment of free-hand transperineal targeted prostate biopsy using multiparametric magnetic resonance imaging-transrectal ultrasound fusion in Chinese men with prior negative biopsy and elevated prostate-specific antigen. BMC Urol. 2017;17(1):52.

24.	Ma TM, Tosoian JJ, Schaeffer EM, et al. The Role of Multiparametric Magnetic Resonance Imaging/Ultrasound Fusion Biopsy in Active Surveillance. Eur Urol. 2017;71(2):174-80.

25.	Recabal P, Assel M, Sjoberg DD, et al. The Efficacy of Multiparametric Magnetic Resonance Imaging and Magnetic Resonance Imaging Targeted Biopsy in Risk Classification for Patients with Prostate Cancer on Active Surveillance. J Urol. 2016;196(2):374-81.

26.	DiBianco J, Mullins J, Allaway M. Ultrasound Guided, Freehand Transperineal Prostate Biopsy: An Alternative to the Transrectal Approach. Urology Practice. 2016;3(2):140.