Theoretical analysis of the oscillating circular piston positive displacement flowmeter: I - modelling the forces acting on the piston

Charlotte E Morton, Ian M Hutchings, John A Williams & Roger C Baker*

Institute for Manufacturing, Department of Engineering, University of Cambridge, 17 Charles Babbage Road, Cambridge CB3 0FS

Abstract

This first part of a three-part paper describes the theoretical modelling of the behaviour of an oscillating circular piston positive displacement flowmeter (also known as the rotary piston flowmeter), by considering the principal forces acting on the piston. These include pressure forces, unsteady flow effects, inlet and outlet flow losses, retarding forces due to thin films and frictional effects. Effects which are too small to affect significantly the dynamic behaviour are not included nor are detailed discussions of liquid leakage. The numerical model lends itself to a time marching solution which applies the forces at each position and deduces the new position and velocities. Thus the local size of the various forces is calculated at each point in the motion for the precise position of the piston at that moment.

In the second part of the paper we set out the calculation process and then demonstrate the validity of our model by comparing it with experimental results. The approach used in this work should lend itself to the modelling of the behaviour of other positive displacement flowmeters.

Keywords: positive displacement flowmeter, oscillating circular piston flowmeter, rotary piston flowmeter, theoretical model of piston movement and pressure losses.

Corresponding author: Roger C Baker

E-mail address: rcb29@cam.ac.uk

Nomenclature

A 	internal cross-sectional area of pipe; area of surface

A, A', B, B'. C, C', D, D' 	identify pressure regions in Fig. 8

A1 A2 A3   	duct cross-sectional areas in Fig. 11

Aloss	used in loss model Equation (18)

Bloss 	used in loss model Equation (18)

C	inclination factor

c	clearance between bottom of centre line of piston and bottom of chamber

c1, c2	constants

D 	hydraulic diameter

F 	force

pressure, friction and gravity forces (including buoyancy of piston) in normal, tangential and vertical directions.

fD	Darcy friction factor (=4 f)

fF	Fanning friction factor

h	film thickness

kf	friction parameter

Ia, Iv 	moment of inertia of the piston about axis and about edge

K 	loss coefficient defined in Eq. 14

L	combined length of up- and downstream pipework; ro-ri; depth of squeeze film

pressure area; length of lubrication holes

l	height of the piston (either internal or external); length of the top rim of the

piston

M	momentum of liquid; total moment on piston

m 	mass of the piston

N	normal force

p	pressure

pi	pressure at inlet to the chamber

po	pressure at outlet from the chamber

pa pb pc pd	pressures identified in Fig. 10

Δp	pressure difference

Re	Reynolds number

Rq	rms roughness

r	radial coordinate, effective radius of curvature

ri	inner radius of piston skirt

ro	outer radius of piston skirt

rp, rp, rp radial position, velocity and acceleration of the centre of the circular piston

t	time

U	velocity of liquid in pipework.

u, v	relative velocity of surfaces in the (x, y) directions

w	relative velocity of surfaces in the z direction given by

x,y,z	Cartesian coordinates

α	is the angle between the contact point of piston/chamber and the radial co-ordinate

β	angular rotation of the piston as defined in Fig. 3

angular rotational acceleration of the piston.

torques due to pressure, friction and gravity

γ	tilting angle of rotor

tilting angular acceleration

η 	coefficient of friction

,, angular position (Fig. 3), velocity and acceleration of the circular piston

Λ	film thickness parameter

ρ	density of liquid

μ	viscosity of liquid

φ 	polar coordinate around piston

Other subscripts

, b 	relates velocities etc to two surfaces within the meter

max, min	maximum and minimum
1.	Introduction

1.1	Oscillating piston flowmeter

The oscillating piston flowmeter (also known as the rotary piston flowmeter) is a type of positive displacement (PD) flowmeter, used to measure the volumetric flow rate of a wide range of clean or well filtered liquids including water, hydrocarbons, other chemicals and liquid foods. The meter  measures the flow rate of a continuous stream of liquid by momentarily trapping a small volume of the liquid in a chamber of known volume and subsequently releasing it on the outlet side. The flow rate is obtained by counting the number of entrapments in a known period of time. This oscillating piston flowmeter differs from virtually all other designs in having a bearingless piston held in position by contact with the measuring chamber. This paper describes the analysis required to model the movement of the oscillating piston together with the flow losses, so that the performance of these flowmeters can be predicted.

PD meters absorb a small amount of energy from the flowing liquid to overcome the viscous drag and the friction at the sliding contacts which results in a pressure loss across the meter. This differential pressure provides the driving force for the meter. Baker [1] described the factors which could affect the volumetric flow rate for a fixed frequency of oscillation: temperature, pressure, viscosity, wear, solid deposits and gas in the liquid. Wear on meter surfaces can cause the displacement volume to increase, and decrease the long term repeatability of the meter, for instance due to leakage. In contrast deposits on the meter surfaces will decrease the displacement volume.  They could also stop the meter from operating.

Viscous losses and turbulence of the flow can contribute to the pressure loss across the meter [2]. The "fill-up characteristic" [3] for the measuring chamber is a function of the angle of rotation, and is generally not linear; the angular velocity of the piston will, therefore, vary. Baker & Morris [4] gave a simple theoretical explanation of some features of the operation of larger types of PD meter.

Since its inception in 1884 the oscillating piston flowmeter has become one of the most widely used positive displacement flowmeters. These meters have wide potential for use. Typical applications range from domestic water metering to high precision, high pressure, measurement of flow injection in oil pipelines and, as well as use with water and chemicals, these meters may be made in designs suitable for liquid foodstuffs and food ingredients.  It consists of a circular piston which moves in a cylindrical measuring chamber that incorporates a cylindrical hub, pin and radial partition (also known as the web) as illustrated in Fig. 1. The centre of the piston, the peg, is constrained to move within the annulus between the hub and pin. The top of the piston consists of an annulus of thickness equal to the thickness of the piston skirt with a recess in the centre of the piston of depth of order 0.5 mm. Within the recess, there are small “lubrication” holes with diameter of order 1 mm, which are equally spaced at half the radius of the piston, and the skirt may have a small cut-out not shown in the figure.

The orientation of the piston and chamber in the model and in our tests was with their axes, hub and pin, vertical. The meter orientation can vary from this, and a similar analysis could be developed for other orientations. The piston profile allows the piston to slide along the partition which also prevents the piston from rotating. Liquid enters the measuring chamber through the inlet port and this causes the piston to move. The oscillation of the piston carries liquid through the meter in essentially self-contained compartments to the outlet port. Fig. 2 illustrates the flowmeter mechanism and the angle θ, used to describe the position of the piston in the chamber.

We reviewed the literature in Morton et al [5] and will, therefore, consider only previous research directly relevant to the analysis. Cloran [6] reported that pressure losses in the oscillating piston meter were likely to arise from the restrictions at the inlet and outlet. Baker [1] confirmed this but also noted that at low flow rates the main cause of pressure loss was the work needed to overcome the internal friction.

Pin

Measuring chamber

Outlet

Inlet

Partition (web)

Hub

(b)

(a)

(c)

(d)

Peg

Circular piston skirt

Profile

Fig. 1 Diagram of the oscillating piston flowmeter (a) measuring chamber (b) circular piston from above (c) circular piston from side (d) circular piston and measuring chamber.

Parry [7] presented a model for the meter using computational fluid dynamics (CFD). The dynamic interaction between the moving parts and the flow in the flowmeter was outlined. The movement of the piston was compared with the movement of a slider-crank mechanism.

The authors have presented data on these meters [5, 8, 9], and these, with other data, will be used to demonstrate the validity of the modelling process in the second part of this paper.

1.2 	Movement of the piston

The piston is free to move horizontally in the measuring chamber, within the constraints of the peg moving in the hub annulus, and the rotation being constrained by the partition. We have chosen to refer to the circular piston although it is often referred to as the rotor. Smaller radial and vertical movements are possible due to the clearances between the piston and measuring chamber. Thus the centre of the piston (the peg) rotates around the chamber pin within the hub and the piston profile slides back and forth along the partition. Fig. 2 illustrates the movement of the piston showing the piston at four positions defined by θ, the angular position of the piston peg around the centre of the chamber, as defined in Fig. 3. Liquid is metered by both the volume outside the piston (outer chamber) and also the volume within the piston (inner chamber). At top dead centre (TDC) or θ =0, both the inlet and outlet ports are closed to the outer chamber. Liquid, however, enters the inner chamber through the inlet port causing the piston to move. This movement expels liquid from the inner chamber through the outlet port, and allows the outer chamber to discharge through the outlet, while fresh liquid is drawn into both inner and outer chambers from the inlet. Incoming and outgoing liquid is now both inside and outside the piston. As the piston moves towards bottom dead centre (BDC) or θ = π, both the inlet and outlet ports are fully open to the outer measuring chamber

(a) θ = 0 (TDC)

(b) θ = π/2

(c) θ = π (BDC)

(d) θ = 3π/2

Fig. 2 Position of the circular piston at four angular positions to illustrate movement of the circular piston around the measuring chamber.

and become closed to the inner chamber. The liquid inside the piston is neutral, but becomes outgoing liquid as the piston proceeds.

When the position of the piston, θ, is combined with the constraint due to the partition, the angular rotation of the piston, β is obtained (Fig. 3). In this figure β is the angle between the radial line of the piston which intersects the partition and the radius of the measuring chamber through the partition, φ is the angle around the piston and α is the angle between the contact point of piston/chamber and the radial co-ordinate of the point at which an element of the pressure is contacting the piston surface.

As the piston moves different points on the piston take different trajectories, the piston movement being governed by the direction of the centre of the piston and the constraint imposed by the partition, Fig. 4. For instance the centre of the piston is seen to move in a circular path, whereas the profile moves close to the partition and the parts of the piston away from the profile move in oval paths. There is, also, a variation in the velocity at different positions on the piston ranging from zero to twice the velocity of the centre dependent on the angular position of the piston. This variation is illustrated in Fig. 5 assuming that the centre of the piston has a constant angular velocity of 10 rad/s. As will be shown, the angular velocity, when in operation, is not constant.

Fig.3 Angular rotation and angular position of the circular piston

Partition/web

(profile motion)

Measuring chamber (contact position)

Pin

Hub

Motion of centre of circular piston

Motion around partition (web)

Fig. 4 Movement of different positions on the circular piston as obtained from the geometrical time stepping program. The colours are used to distinguish one curve from another.

0                             π/2                          π                             3π/2                        2π

profile

piston

piston skirt

Fig. 5 Velocity of different points on the circular piston assuming a constant angular velocity.

An initial stage in the development of the numerical model was the development of a geometric model which enabled Figs. 4 and 5 to be obtained, and formed the basis for the development of the dynamic model, which used a time-marching procedure to follow the variation of all parameters including the speed of oscillation. The contact position on the piston is the point adjacent to the chamber wall. This paper describes the development of the dynamic oscillating piston flowmeter model.

We shall first discuss the equations of motion for the dynamic model, and will then proceed to describe the two sub-models: the pressure force model and the friction model.

2.	Equations of motion

2.1	Tangential and radial motion

Simplified equations of motion for the piston can be derived by treating its centre as a

particle [10] whose position is described by polar coordinates (rp ,θ ) relative to the centre of

the pin. Thus we may write

(1)

(2)

where m is the mass of the piston, FT  and FN  are the resultant tangential and radial forces acting upon it. In practice the first term in Eq. (1) dominates since rpθ≫rpθ due to the small clearance between the piston and chamber wall, so that it can be simplified to:

(3)

This enables the equation of motion in the tangential direction of motion to be solved, before solving the equation in the radial direction.

Thus the variation in radius of the centre of the piston is limited by the design clearances. For the intermediate size of meter, with an average value of this radius of about 3.75mm, the average tolerance was about ±0.02mm. We have, therefore, taken the motion of the piston centre, for Equation (1), to be in the θ direction. While the radial movement is very small, it is important in the context of radial force and friction.

There will also be a moment or torque,  on the piston. This torque arising from the differential pressures and from contact forces with the partition causes the piston to rotate until prevented from doing so by the partition. We may write

(4)

where Ia is the moment of inertia of the piston about its axis and  is the angular rotational acceleration of the piston. Morton [11] gives a fuller description of the analysis.

2.2	Vertical motion

The vertical force and torque are given in Eqs. (5) and (6) respectively.

(5)

(6)

where is the acceleration in the vertical direction (z-axis), is the moment of inertia of the piston about a tangent to its skirt and is the tilting angular acceleration. Integrating Eqs. (5) and (6) twice enables z and γ to be found.

In order to solve the dynamic movement of the piston, we proceed to obtain the forces and moments: . These forces are due to pressure, friction and gravity (including buoyancy of the piston).

2.3 Mass and inertia of piston

In considering the motion of the piston we should note that the piston is surrounded by and filled with liquid. If the liquid within the piston was unchanging relative to it, then the mass of that liquid would simply need to be added to that of the piston, and the forces would act on the combined mass. Additionally, at least some of the liquid outside the piston may move with it. For the densest liquid used the mass within the piston was about 2.5 g, while the pistons themselves had masses of 3.6 g, 6 g and 24 g for the PMMA, carbon (carbon graphite)  and SS pistons respectively.

In practice, there are only two points in the oscillation of the piston where the liquid inside or outside the piston is trapped relative to the piston; these are momentarily at TDC for the outside of the piston and at BDC for the inside. In all other positions the liquid within or outside the piston is in relative motion, either from the inlet duct or to the outlet duct. The flow rate through the chamber will be at a velocity greater than that of the oscillational movement, and the moving liquid thus only momentarily applies pressure to the piston where it impinges on its surface.  To account fully for these aspects of the liquid motion would require analysis of the dynamics of the complex flow by CFD. In the present model these effects have been ignored and the mass of the piston alone has been used in the equations of motion; the validity of this approximation is demonstrated in Part II of this paper by comparing the predictions of the model with the experimental measurements for the pistons of different density. The experimental data suggests that the added mass is a factor which will in reality be of small effect.

3.	 Pressure forces creating horizontal motion

Pressure forces result from:

- external pressure applied to the meter due to its position in the pipeline;

- lubrication type pressure differences due to flow through converging gaps between solid surfaces driven by the external pressures;

- squeeze films due to solid surfaces approaching or moving away from each other.

These are illustrated in Fig. 6.

(b)

x

z

x

z

h0

z

pa

(a)

h0

ps

x

u

h0

(c)

Fig. 6 Mechanisms for pressure development within the flowmeter: (a) externally generated pressure; (b) converging channel, (c) squeeze film. (Hamrock [12])
3.1	External pressure

The dominant forces due to pressure are created by the pressure difference between the liquid at inlet and outlet. At θ = 0 (TDC) the liquid in the outer measuring chamber is not connected to either the inlet or outlet. This is also the case for the liquid inside the piston at θ = π (BDC). In each case, for the purposes of modelling, the pressure of the trapped liquid is assumed to be the mean of the inlet and outlet pressures. It is assumed that liquid within the annulus of the hub is connected to the inlet for angular positions of θ > 0 and θ < π and connected to the outlet between θ > π and θ < 2π.  This is because the profile of the piston allows liquid to enter the annulus freely from either the inlet or outlet depending on the angular position of the piston.  Fig. 7 shows four positions of the piston and the assumed pressure of the liquid within the flowmeter at each of these positions.

This pressure force can be resolved parallel and perpendicular to the direction of motion, as illustrated in Fig. 8 which shows the piston at θ = π/2.  Parallel to the direction of motion some of the pressure forces cancel with one another and others do not. For example when θ = π/2 (Fig. 8(c)) the pressure forces A and B cancel as do C' and D'. Pressure forces from the outside of the piston do not cancel with those from the inside of the piston due to the difference in dimensions. The remaining pressure forces, A', B', C and D (as shown in Fig. 8(d)) can then be added together.

(d) θ = 3π / 2

(c) θ = π  BDC

(a) θ = 0 TDC

(b) θ = π /2

KEY

= Pi

= Po

= Pneutral

Fig. 7 Pressure of the liquid within the flowmeter at four different positions

A

B

B’

D

D’

C’

(a)

(c)

(e)

(b)

(d)

(f)

A’

C

Fig. 8 Illustration of pressure forces with the piston at θ = π/2; (a) acting perpendicular to the surface of the piston; (b) piston separated into four internal and four external quadrants; (c) resolved in direction of motion; (d) resolved in direction of motion with cancellations;  (e) resolved normal to the direction of motion; (f) resolved normal to the direction of motion with cancellations.

0                           π/2                        π                          3π/2                      2π

(a)

0                           π/2                        π                          3π/2                      2π

(b)

Fig. 9 Variation in the force on the piston due to the differential pressure  (a) Variation in the cross-sectional area over which the differential pressure acts in the tangential direction to motion (b) Variation in the cross-sectional area over which the pressure force acts in the direction normal to the motion.

Similarly, if the pressure forces are resolved in the direction normal to motion, some of the pressure forces cancel. For example when θ = π/2 (Fig. 8(e)) the pressure forces B and C cancel with each other as do B' and C'. Pressure forces A and A' and D and D' cannot be cancelled with one another because of the difference in the inner and outer dimensions of the piston.  The remaining pressure forces, A, A', D and D' (Fig. 8(f)) can be added together and result in a normal force acting to reduce the radius of rotation, rp of the piston. Although these normal and tangential forces are strictly related to the centre of the piston, the alignment of the centre of the piston will be such that they can be considered as forces in the polar r and ϴ directions as defined in Figure 3.

The code for the model has been set up to calculate both the internal and external pressure forces and their directions.

The tangential pressure force is the integral of the force due to the pressure, pα, resolved in the tangential direction of motion (both for internal and external pressure) as illustrated in Figure 8 for one position of the piston. The integral is expressed as

FT=l02πp(α)rsin d			                                       (7)

where

=θ-(φ-)

and where l is the height of the piston (either internal or external), r is the radius of the piston (either internal or external), , θ,   and φ are shown in Figure 3.

In our initial model we used the approximation

=θ-φ

Our calculations for comparison with the experimental data made use of this approximation. This approximation introduced a small error for the average pressure loss, for example 1.5% for a small size of meter for liquid of 10 mPas, and also an increase in the predicted pressure loss across the piston of 7-8%.

The tangential force is essentially the product of the differential pressure across the piston and the cross-sectional area of the piston. At TDC the differential pressure acts over the internal dimensions of the piston and is a minimum. At BDC the differential pressure acts over the external dimensions of the piston and is a maximum. At θ = π/2 (Fig. 8(a)), the cross-sectional area falls between the two. The variation in the cross-sectional area is shown in Fig. 9(a).

Similarly, the normal pressure force can be found by integrating the pressure force in the direction normal to the motion (both for internal and external pressure see Figure 8)

FN=l02πp(α)rcos d	                     				    (8)

The arrow in Figure 3 shows the positive angular direction for α which will apply in Equations (7) & (8) for external pressures on the piston.  For internal pressures the signs will be changed to allow for radial and tangential forces in the opposite directions. Both the cross-sectional area and direction over which the resolved pressure forces act in the normal direction vary during the rotation. This variation in the cross-sectional area is displayed in Fig. 9(b).

At both θ = 0 (TDC) and θ = π (BDC), the normal pressure forces around the piston are balanced. When 0 < θ < π, the effect of the resultant normal pressure force is to reduce the radius of rotation (as shown in Fig. 8(f)). When  π < θ <2π, the net normal pressure force acts to increase the radius of rotation.

In considering the action of the pressure forces on the piston, two factors can be considered which will modify the effective applied pressure. These are:

- the effect of unsteady flow in the inlet and outlet pipes;

- the losses which occur as the liquid flows from the inlet pipe into the chamber, and from the chamber into the exit pipe.

3.2	Unsteady flow effects - Pressure transient model

If the flow rate through the flowmeter is not constant, the velocity of the liquid in the adjacent pipework will either be increasing or decreasing. Since the liquid in the pipework prior to and after the flowmeter possesses mass and thus momentum, these changes in velocity will generate differential pressures, as described by Chadwick et al [13]

Suppose, mass of liquid in pipework 			    	    (9)

then, momentum of liquid in pipework 		    (10)

where ρ is the (assumed constant) density of the liquid, A is the internal cross-sectional area of the pipe, L is the combined effective length of the up- and downstream pipework and U is the velocity of the liquid. If the flow is retarded at the flowmeter, the liquid undergoes a rate of change of momentum given by

(11)

This change in momentum requires a force, F, which exists during the deceleration and acceleration of the liquid. This force causes an instantaneous change in pressure, Δp, defined in Eqs. (12) and (13).

(12)

(13)

The pressure difference caused by the unsteady flow is therefore affected by the density and the acceleration or deceleration of the liquid in the pipe and the length of the pipework. A change in any one of these factors can change the pressure variation observed across the flowmeter.

There is also a pressure loss due to liquid friction as liquids flow through the pipework before and after the flowmeter. The general expression for the pressure loss can be written as

(14)

where K is the loss coefficient. For a straight pipe K takes the form

(15)

where fD is the Darcy friction factor, L is the combined length of up- and downstream pipework  and D is the hydraulic diameter. The Darcy friction factor, fD, is equal to four times the Fanning friction factor, fF.  The value of the friction factor, fD, is dependent on the Reynolds number. Provided the flow is laminar we may write

(16)

Unless the length of pipework is considerable or the liquid is very viscous, the pressure losses from the liquid friction are small.

A schematic diagram of the flow rig for testing the flowmeter is given in Fig. 10. More complete details are given by Morton [11]. The change in pressure difference across the flowmeter (the difference in pressure between pb and pc in Fig. 10) is given by

(17)

where g is the acceleration due to gravity and h is defined in Figure 10. This difference in pressure across the flowmeter is dependent on the length, diameter and friction factor of the pipework, the density, velocity and acceleration (or deceleration) of the liquid being metered and the pressure difference between the header and sump tank. The flow will be assumed to be incompressible and the pipe material rigid.

header tank

pa

flowmeter

h

sump tank

pd

pbb

pc

Fig. 10 Schematic diagram of the flow rig for testing the flowmeter

3.3	Pressure losses in inlet and outlet ducts

There is a loss of total pressure due to liquid friction as liquid flows through a uniform pipe. Additional losses occur due to rapid changes in liquid velocity and direction which occur at bends and changes in pipe cross-section such as contractions and expansions. The pressure loss coefficient, K, in turbulent flow may in some cases be independent of the Reynolds number, but in laminar flow may be inversely proportional to the Reynolds number. Turbulence typically occurs in pipe flow at Reynolds numbers of greater than about 2000. However when there are abrupt changes in area and direction in components, turbulence may be induced at lower Reynolds numbers.

Flow enters the flowmeter through the inlet port and exits through the outlet port. A two-dimensional diagram of the entry and exit ducts is given in Fig. 11. The flow path can be summarised as a sequence of four 90° bends (or right angled pipe junctions) with associated changes in section joined by straight pipe sections and the measuring chamber, as set out in Fig. 11 and Table 1. Data for loss coefficients for the duct in Fig. 11 were not available, and so best estimates were required based on available data. For this purpose data was obtained from Miller [14]. Data on loss coefficients for the inlet and outlet ducts were considered based on: sudden contractions and expansions, re-entrant inlets, T-junctions with one branch closed, mitre bends and also the possibility of swirl arising from the offset between the entry pipe and the inlet duct. Flow losses in the chamber without a piston in place are difficult to estimate because of the complex nature of the flow. The ranges are given in Table 1. Other flow losses due to friction through the other straight pipes from Eq. (15) caused typical loss coefficients of less than 0.1.

Experimental values of the loss coefficients were taken when the piston was omitted from the measuring chamber as tests were carried out with this arrangement which enabled the pressure losses to be monitored. Due to the problem of modelling the losses, the experimental value of losses was used to develop a simple model as given in Equation (18). This is based on an assumed laminar/turbulent relationship [14] and is given in Fig. 12

K=Aloss+BlossRe							     (18)

where Aloss=5.94 with Bloss is in the range 800 to 1,000. The relationship is compared with the data in Fig. 12 for two values of Bloss.  Bloss=800 was used in the comparison between the model and the experimental data. Aloss is essentially the value of the loss coefficient, K, for large values of Reynolds number in this model and Bloss is related to the lower Reynolds number dependent region.

Entry pipe

Exit pipe

A1

177mm2

A1

177 mm2

A2

28 mm2

A3

90 mm2

A2

28 mm2

Flowmeter chamber

Inlet duct

Outlet duct

Fig. 11 Schematic two-dimensional diagram of the entry and exit ducts

Table 1 Loss coefficients for flow through meter with no piston based on the

velocity in the inlet and outlet ducts

| Flow path | Loss coefficient for components adjusted to velocity in A2 |
| --- | --- |
| contraction from A1 to A2 | 0.3 to 2.0 |
| expansion from A2 to A3 | 0.9 to 1.0 |
| contraction from A3 to A2 | 0.5 to 2.0 |
| expansion from A2 to A1 | 0.5 to 1.0 |
| Flow through inlet and outlet ducts with allowance for swirl | 0.4 to 0.6 |
| Flow through 180o bend in chamber including loss of KE | 0.5 to 1.2 |
| TOTAL | 3.1 to 7.8 |

Bloss

Bloss

Fig. 12 Loss coefficient for all Reynolds numbers

3.4	Retarding forces due to pressures in narrow passages

The important flows through the narrow gaps around the piston can be investigated using Reynolds equation [12] which provides a relationship between hydrodynamic pressure gradients, the geometry of the surfaces, their relative sliding velocity and the properties of the liquid. Reynolds equation can be written as

Squeeze

Local expansion

Poiseuille

Couette

(19)

where u, v and w are the velocities of surfaces in the x, y and z Cartesian directions respectively, and subscripts a and b refer to the two surfaces, h is the film thickness in the z direction as a function of x and y. The x and z axes are shown in Fig. 6 and the y axis is perpendicular to both the x and z axes.

The Poiseuille and Couette terms describe respectively the net flows due to pressure gradients and shear driven by solid surface velocities within the lubricated area. The squeeze term describes the flow due to squeezing motion (as shown in Fig. 6(c)) and the local expansion term describes the flow resulting from any change of density [12].

3.5	Squeeze films

Squeeze films can be generated when two surfaces approach one another and the liquid film acts as a cushion between the two surfaces such as between: the outside of the piston skirt and the measuring chamber, the inside of the piston skirt and the hub, the peg and the hub and the peg and the pin. When the surfaces move apart the reverse effect occurs and reduced pressures are generated which, potentially, may lead to cavitation. This does not occur in the flowmeter because the mean pressures are high enough to avoid these conditions.  Movement in the radial direction is illustrated in Fig. 14.

While liquid density, ρ, and viscosity, μ, can be functions of pressure and temperature, these effects are small within the flowmeter and we can assume that density and viscosity are constant. Assuming that , that the motion is restricted to the squeeze velocity,, (omitting the subscript b) and that liquid flow is restricted to the x direction, Eq. (19) becomes

(20)

Integrating once gives

(21)

(a)

(b)

h0

rc

ror

h0

t

h

w









rpeg

rpin

x

x

Fig. 14 Movement in the radial direction (a) between piston skirt and measuring chamber (conformal contact) (b) between peg and pin (non-conformal contact)

The boundary condition atis, giving c1 = 0.  Integrating Eq. (21) gives a pressure distribution

(22)

Pressure is localised near the line of centres, i.e. where x is small, and in this region the film thickness, h, can be approximated by the expression

(23)

(Neglecting further terms in this binomial expansion introduces an error of order 0.7% or less into the integration.)

In Eq. (23) r is the effective radius of curvature and is given by the expression  in which r1 and r2 are the radii of the two surfaces [15]. The derivative of the film thickness is thus

(24)

and substitution of Eq. (24) into Eq. (22) gives

(25)

The pressure will be zero when h is large, giving c2 = 0. The pressure distribution is thus

(26)

The pressure force acting in the radial direction due to the squeeze effect is obtained by integrating Eq. (26) to give

(27)

where L is the depth of contact in the y direction. Since the pressure distribution is symmetric it generates no net force in the direction tangential to motion.

### 3.6	Pressure forces due to trapped liquid at TDC and BDC

There is a maximum theoretical area for the inlet and outlet ports, such that at no point in the oscillation of the piston are both in direct communication with each other or both closed. The actual areas of the inlet and outlet ports are less than these theoretical maxima. A schematic diagram of the actual and theoretical maximum sizes for the inlet and outlet is shown in Fig. 15.  This means that for certain positions during the oscillation there are volumes of liquid which have no means either to enter or leave the measuring chamber. This in turn causes either a high or low pressure which creates a decelerating force. As the piston approaches TDC, as in Fig. 16 (a), the liquid in the chamber (Co), has no means of leaving the measuring chamber except by leakage between the piston and chamber. . The stepping program calculates, from the geometry, the leakage and the resultant force which builds up in the trap, and balances this with the other forces on the piston. Thus, as the piston moves towards TDC, pressure within Co builds up so retarding the piston. Similarly, as the piston moves away from TDC, Fig. 16 (b), liquid has no means to enter the chamber (Ci), and a reduction in pressure is created within Ci resulting in the piston experiencing a decelerating force. Similarly, as the piston moves through BDC decelerating forces are created just prior to (Po) and after (Pi) BDC as illustrated in figs. 16(c) and 16(d).

The total area of both the inlet and outlet ports can be plotted as the piston peg rotates within the hub and this is illustrated in Fig. 17.

Theoretical maximum area of outlet port

Theoretical maximum area of inlet port

Actual area of outlet port

Actual area of inlet port

Fig. 15 Schematic diagram of the theoretical maximum and actual areas of the inlet and outlet ports

Po

Pi

(c)

(d)

Co

Ci

(a)

(b)

Fig. 16 Schematic diagram of the flowmeter illustrating the trapped liquid: a) as piston approaches TDC, b) as piston departs TDC, c) as piston approaches BDC, d) as piston departs BDC

0                          π/2                         π                          3π/2                       2π

(a)

0                          π/2                          π                          3π/2                       2π

(b)

Piston

Piston

Fig. 17 a) area of inlet port open to the outer measuring chamber (Chamber) and to the inner measuring chamber (Piston), b) area of outlet port open to the outer measuring chamber (Chamber) and to the inner measuring chamber (Piston),

Fig. 17(a) shows the area of the inlet open to both the outer measuring chamber and to the inner measuring chamber (within the piston). The area of the inlet open to the outer measuring chamber is zero between θ = 0 and θ = π/4. The volume of the outer measuring chamber starts to increase from θ = 0 and the liquid has no means of entry to the chamber except via the small cut-out in the piston skirt and through leakage. Similarly, the area of the inlet open to the inner measuring chamber (the piston) is zero between about θ = 2π/3 and θ = 4π/3. Therefore, between θ = π and θ = 4π/3 liquid can enter the piston only through the cut-out in the piston skirt and through leakage.  Similarly, Fig. 17(b) shows the area of the outlet open to both the outer measuring chamber and the inner measuring chamber. This means that between θ = 0 and θ = π/4, θ = 2π/3 and θ = 4π/3, and between θ = 7π/4 and θ = 2π, a decelerating force will act on the piston.

The pressure loss caused by the trapped liquid is calculated by estimating the losses caused as the liquid passes through a narrow cut-out on the bottom of the piston. Eq. (14) allows the pressure drop to be calculated provided we can estimate the loss coefficient, K.

Allowing for the sudden expansion and contraction and the area ratios for the contraction and expansion [14] we estimated a value of K 1. When the area of the outlet port reduces to less than the area of the cut-out, liquid will start to flow through the cut-out until the outlet is fully closed. The flow rate through the cut-out is therefore equal to the rate of change in volume of the measuring chamber and this is calculated by the geometric model. The pressure loss due to trapped liquid can therefore be calculated and is used to calculate the tangential and radial pressure forces on the piston. The volume of trapped liquid just before and just after BDC, is very much less than that at TDC due to the smaller radius of the hub and the reduced internal height of the piston. This means that both the pressure loss and the decelerating force around BDC are very much smaller than at TDC.

4.	 Pressure forces in the vertical direction

4.1		Externally generated pressures in the vertical direction - from below the piston

In addition to forces generated in the horizontal or x-y plane, the piston experiences

pressures on its horizontal faces and so forces in the vertical direction. In estimating these it

is assumed that

- the pressure inside the piston will be either inlet or outlet pressure depending on whether the liquid is connected to the inlet or the outlet (except for a momentary neutral position at BDC).

- the pressure inside the hub is assumed to be equal to inlet pressure between θ = 0 and θ = π, and outlet pressure between θ = π and θ = 2π

- the pressures above the hub and below the piston skirt can be calculated using Reynolds equation (19), converted to radial coordinates, in the form

(28)

hmin

r

hmax

h

γ

w

h=c on centre line

Fig. 18 Schematic diagram of piston tilting with axial and angular movement

Initially, it is assumed that the film thickness, h, does not vary in the radial or angular position (i.e. the piston is parallel to the measuring chamber). Eq. (28) can be integrated twice and combined with the boundary conditions at  and at to give the pressure distribution

(29)

The total pressure force from the bottom rim of the piston is the integral of the pressure over the area of the piston

(30)

where r and φ are the radial and angular coordinates in cylindrical polar coordinates related to the piston. The effect of tilt was estimated to be negligible [11, 12].

4.2		Externally generated pressures in the vertical direction: from above the piston without lubrication holes

In the particular meter investigated, the top of the piston consisted of an annulus of thickness equal to the thickness of the piston skirt with a recess in the centre of the piston of depth 0.5 mm. Within the recess, there were seven lubrication holes of diameter 1 mm, equally spaced at half the radius of the piston, as illustrated in Fig. 19.

The lubrication holes will affect the liquid pressure above the piston. Initially, the pressure was calculated with no lubrication holes. The pressure force was then modified to include the lubrication holes. With no lubrication holes, the pressure within the recess was assumed to be the mean of the inlet and outlet pressures. The pressure distribution around the rim of the piston was obtained using Reynolds equation, Eq. (19).

(b)

Top rim of piston circular piston

Lubrication hole

(a)

Top rim of piston circular piston

Lubrication hole

Fig. 19 Diagram to indicate the position of lubrication holes (a) Top view of the piston, (b) sectional view of the piston

The pressure distribution obtained was

- for and for all		                             (31)

- for and                               (32)

- for and                              (33)

The pressure is equal to the mean pressure of the inlet and outlet within the recess in the top of the piston for all angular positions. The pressure over the top edge of the piston varies linearly between the mean pressure and either the inlet or outlet pressure depending on both the angular position of the piston and the position on the top of the piston. The approximate pressure distribution expressed by equations 31-33 is considered to be adequate for this calculation and, as will be shown in part II of this series of papers, the agreement between theory and experiment is within about 5%.

4.3	Externally generated pressure forces from above the piston with lubrication holes

With lubrication holes, the pressure in the recess above the piston cannot be assumed to be the mean of the inlet and outlet pressures. There will be flow through the holes dependent on pressure difference, Δp, across the holes. For one hole this is

(34)

where a is the radius and L is the length of a lubrication hole.

For angular positions between θ = π/4 and θ = π, the liquid in the piston below all the lubrication holes is connected to the inlet. This includes the liquid between the hub and pin because of flow around the profile of the piston. Therefore, the liquid in the piston below all lubrication holes has a pressure equal to the inlet pressure. This is illustrated in Fig. 20 with θ = π/2.

Any liquid that flows through the lubrication holes will therefore have to flow over the rim of the piston and the flow through the seven lubrication holes must be equal to this. The flow rate through the lubrication holes is dependent on the difference in pressure between the inlet and recess pressures. The flow rate over the rim is dependent on the difference in pressure between the recess and outlet. The recess pressure adjusts such that the flow through the lubrication holes equals the flow over the rim, assuming that this is due to the pressure difference only, and is given by

(35)

where l is the length around the top rim of the piston and , as shown in Fig. 20.

l

L

Fig. 20 Positioning of the lubrication holes at θ = π/2

Combining Eqs. (34) and (35)

(36)

from which the pressure in the recess, pr can be obtained. The pressure difference across the lubrication holes is reduced in order to balance this flow.

Between θ = π and θ = 7π/4, the liquid below all lubrication holes has a pressure equal to the outlet pressure. A similar expression to Eq. (36) is obtained (Eq. 37) and the pressure within the recess, pr can be found.

(37)

For angular positions around TDC (i.e. between θ = 0 and θ = π/4 and between θ = 7π/4 and  θ = 2π), there are lubrication holes over both the inlet and outlet. Therefore flow can take place between the lubrication holes. At θ = 0 (TDC), there are two lubrication holes above the inlet and two above the outlet. It is assumed that there is no flow over the top rim. Therefore the flow through the holes in the inlet is equal to the flow through the holes in the outlet; hence the pressure in the recess is assumed to be equal to the mean pressure of the inlet and outlet.

As the piston moves around towards θ = π/4 six lubrication holes are open to the inlet and one to the outlet. The pressure within the recess, pr, is obtained using

(38)

where a and L are the radius and length of the lubrication holes respectively and l is the length of the top rim of the piston (as shown in Fig. 20).

The pressure within the recess on the top of the piston, pr, can be obtained from Eqs. (36)–(38). The pressure distribution over the top of the piston can be assumed to be

- for  and for all			                                         (39)

- for and                                      (40)

- for and                                     (41)

The pressure force above the piston can be obtained from the pressure distribution in Eq. (32). The sum of the pressure forces from above and below, added together, gives the combined pressure force. Between θ = 0 and θ = π, the pressure force is in the upwards direction and acts to separate the bottom of the piston and measuring chamber. Between θ = π and θ = 2π, the pressure force is in the downwards direction and acts to separate the top of the piston and measuring chamber.

The tilting moment due to an externally generated pressure is

(42)

where r and φ are the radial and angular coordinates of the piston in cylindrical polar coordinates. The angle through which this moment acts will vary with angular position. The integration was performed numerically using MATLAB and is calculated for arbitrary angles between 0 and 2π. Fig. 21 illustrates the variation in the angle of tilting during the rotation.  The tilting moment is a maximum at TDC (θ = 0) and a minimum at BDC  (θ = π). At TDC (θ = 0), the tilting moment acts perpendicularly to the angular position of the piston (φ = π/2). As the piston moves around the chamber, the axis about which the tilting moment acts also moves as shown in Fig. 21. If the piston sits parallel to the bottom of the measuring chamber, there are no convergent (or divergent) channels and no hydrodynamic pressures are generated. If the piston tilts, hydrodynamic pressures will be generated, and the appropriate Reynolds equation in cylindrical polar coordinates is

(43)

4.4	Squeeze Film Effects in the vertical direction

As the piston moves vertically, pressures will be generated due to squeeze film effects above and below the piston and also above the hub on the measuring chamber. These three contact points can all be considered to be thin annuli with an outer radius of ro, a mean radius of rm and an inner radius of ri. We suppose that the film thickness between the surface of the piston and measuring chamber is h and initially that the piston and chamber faces are parallel. Making use of the equivalent of Eq. (21) in cylindrical coordinates and making use of the radial dependence only we obtain the following equation

(d)

(b)

(e)

(c)

(a)

Fig. 21 Variation in angle of tilting during rotation (where     is out of the page and     is into the page); (a) θ = 0; (b) θ = π/3, (c) θ = 2π/3, (d) θ = 4π/3; (e) θ = 5π/3

(44)

Integrating Eq. (46) twice with respect to r gives

(45)

The boundary conditions are that at the inner and outer radii the pressure created by squeeze effects is zero, hence

(46)

The pressure distribution obtained in Eq. (46) can be either positive or negative depending on the sign of the time derivative of h. The force due to squeeze effects can be obtained by integrating Eq. (46) over the area of the rim of the piston, and following Etsion [16-18] we write

(47)

where rm is the mid radius. The torque due to the squeeze film can also be obtained by integrating

(48)

If the piston tilts, the pressure distribution will be affected. The film thickness, h, and its time derivative are defined in Eqs. (49) and (50).

(49)

(50)

where c is the clearance along the centreline of the piston, γ is the tilting angle, w is the axial vertical velocity along the centre line of the piston (Fig. 18) and is the contribution due to the tilting of the piston.

Substituting the derivative of h into Eq. (46) we obtain

(51)

If the tilt is small then the variation in h across the contact point, for instance radially across the rim, may be neglected. Making use of the pressure distribution in Eq. (51) the expression for the squeeze force for small angles of tilt, when the inner and outer radius ratio is approximately unity, was obtained by Etsion [16-18] as

(52)

where c is the mean clearance between the bottom of the skirt of the piston and the chamber. The pressures created between the surfaces of the piston and measuring chamber will also cause a tilting moment. Making use of the pressure distribution in Eq. (51), the tilting moment was obtained [16-18] as

(53)

For a derivation of Eqs. (52) and (55), the reader is referred to Etsion [16, see also 17, 18] who obtained the integral of pressure for squeeze effects in radial face seals, and the expressions for F and M. We have used his approximation for small angles of tilt which introduces an error within about ±0.0002 Nm for values of c above 5.5 μm and will tend to underestimate the squeeze effect for lower values of c. As will be shown in part II of this series of papers, the agreement between theory and experiment for the effect of tilt was within about 5%.,

The squeeze force and moment oppose the motion and the tilting of the piston. The squeeze force is assumed to act at the same position as the externally generated pressures.

4.5	Upward force due to inlet jet

A jet is produced when liquid enters the measuring chamber from the inlet duct. When the piston covers the inlet port, this will create an upward force on the bottom of the piston, as illustrated in Fig. 22. The upward force caused by the jet of liquid can be estimated using the momentum balance so that

(54)

where U is the velocity of the liquid in the inlet duct and A is the area of the inlet duct. The area of the inlet varies during rotation, and is a maximum at TDC (θ = 0) where the piston fully covers the inlet and is minimum at BDC (θ = π), where the area of the inlet open to the piston will be zero. The maximum force from the jet can be calculated at 90 l/hr with the inlet fully open (θ = 0) and is found to be around 0.03 N. This can be compared with typical pressure forces from the inlet and outlet pressures at the same flow rate of about 0.3 N.

The pressure forces created by the jet will also create a tilting moment and this is illustrated in Fig. 23. The tilting moment is

(55)

where r is the lever arm of the force from the centre of the piston. The maximum moment from the jet at a flow rate of 90 l/hr with the inlet fully open has a value of around 0.0003 Nm. There will be some dissipation as the jet emerges into the chamber. At the same flow rate and position, the moment from the externally generated pressures is around 0.002 Nm. The moment from the pressure forces can be added to that caused by the jet. This is done by calculating the turning force at every step of the calculation.

4.6	Weight and buoyancy force

As well as the downward force caused by the weight of the piston, an upward buoyancy force is produced by the liquid in which the piston is immersed, equal to the weight of the liquid displaced by the piston. The resultant downward force on the piston is affected by the density of both the piston and the liquid in which it is immersed.

(a)

Partition

Outlet duct

Inlet duct

Measuring chamber

Circular piston

Sectional view  A

A'

Flow direction

Flow direction

Inlet duct

Outlet duct

Circular piston

Measuring chamber

Partition

Jet creates force on bottom of the circular piston disc

(b)

Fig. 22 Effect of jetting liquid: (a) position of the piston, (b) sectional view A-A'.

r

Fig. 23 Tilting moment caused by inlet jet (where     is out of the page and     is into the page);

5.	Friction losses

If the surfaces of the moving components within the meter come into mechanical contact,

i.e. cease to be separated by a continuous liquid film, then additional frictional losses will

occur. Potential sources of friction involving the piston and chamber are thus at the top and

bottom surfaces of the piston, at both the inside and outside surfaces of the skirt and around

the hub, the partition and the peg. The extent of the frictional interaction depends on the film

thickness and surface roughness of the two surfaces. A mean line of length L can be defined along the surface and the rms surface finish, Rq, will be defined in the standard way as

(56)

where z is the vertical deviation of the surface contour from the mean. The frictional interaction can be thought of as a function of the film thickness parameter, Λ, which is the ratio of the minimum fluid film thickness, ho, to the mean rms surface finish of the two surfaces involved, Rq,a and Rq,b

(57)

A typical value for the rms surface finish of the piston and chamber is 0.8 μm.

The dynamic flowmeter model determines the film thickness between any two surfaces from the forces acting in the meter and determines the film parameter, , whose value distinguishes three different regimes of lubrication.

- Hydrodynamic lubrication (Λ ≥3) is known as full fluid lubrication; the frictional forces are attributable to the viscous losses of the liquid.

- Partial lubrication (1≤ Λ <3) involves some asperity contact although most of the surfaces are separated by a liquid film.

- Boundary lubrication (Λ <1) involves close surface contact where surface asperity contact dominates.

Ideally, opposing surfaces are completely separated by a liquid film (i.e. Λ > 3) so that the frictional force is directly attributable to the viscous losses in the liquid and surface finish is of relatively little importance. The hydrodynamic frictional force for two parallel surfaces (Fig. 24a) can be written as

(58)

where μ is the  liquid viscosity, u is the velocity of the moving surface, L is the length of the surface in the x direction, B is the length in the y direction, h0 is the minimum film thickness and A= LB is the area of the surface. For a tilted surface (Fig. 24b), the minimum film thickness is hmin and the maximum is hmax. The inclination factor, C, can be defined as

(59)

so that the film thickness, h, can, therefore, be written in terms of x as

(60)

With these equations a close approximation to the hydrodynamic frictional force for tilted surfaces may be derived from

(61)

giving the hydrodynamic force for tilted surfaces as

(62)

For effectively unlubricated surfaces in contact the frictional force, F, is expressed conventionally as

(63)

where η is the coefficient of friction and N is the normal force. In the presence of a liquid, frictional resistance, F can be thought of as consisting of two parts: the first due to the asperity interactions, and the second due to the shearing of the liquid film. Hence we can express the friction as

(64)

where Fa,b is due to boundary friction and Ff,H is due to hydrodynamic friction.

(a)

(b)

u

H

L

h

O

x

O

h

γ

u

H

L

x

hmin

hmax

Fig. 24 Geometry of two surfaces (a) parallel surface (b) inclined surface;

As the separation of the two surfaces, measured by the film parameter Λ, increases, so the amount of asperity contact falls. When Λ≥ 3, it is reasonable to assume that no asperity contact takes place. Defining a parameter kf, to describe the partition between solid and lubricated contact such that when Λ≤ 1 the value of kf = 1, for Λ > 3 kf = 0 and over the range 1< Λ <3  its value varies linearly with Λ, then we can express the frictional interaction between lubricated parallel surfaces as

(65)

If the two surfaces are not parallel to one another, the hydrodynamic frictional force will be affected and the frictional force for all film parameters for tilted surfaces can be written as

(66)

where C is the inclination factor defined in Eq. (59).

6. Brief introduction to the calculation method

The dynamic model was solved using MATLAB to apply a time stepping approach. First the input data was defined including the properties of liquid, dimensions of meter and external features. The time step was defined. Each variable was computed at every time step and the results were fed back into the model to calculate the variables at the next time step.

The geometry of the measuring chamber and initial conditions were defined including the initial position of the piston. The dynamic model included forces in the tangential, radial and vertical directions, which resulted in movement of the piston. From the initial conditions, the position of the piston could be calculated, clearances could be computed between the piston and measuring chamber and the frictional force calculated from the clearances.

The main section of the computer model was a loop which enabled values of each variable to be calculated at each time step. The pressure difference across the flowmeter system was defined in the input data. The actual pressure difference across the flowmeter was dependent on the rate of change of flow within the surrounding pipework. Velocity and acceleration of the liquid in the pipework were calculated, and hence the pressure difference across the flowmeter. Pressure loss at inlet and outlet ports was subtracted from the pressure difference across the flowmeter to give the pressure difference across the piston. This pressure difference was used to calculate the pressure forces around the piston in the tangential, radial and vertical directions.

The dynamic equations were solved, and this enabled the new position of the piston and clearances to be found, and the frictional force calculated. The volume of liquid entering and leaving the flowmeter was calculated, and this enabled the ideal flow rate to be calculated.

7.	Discussion

In this part of the paper we have set out the pressures and the resultant forces acting on the piston. We have modelled these as precisely as possible, while making reasonable approximations and neglecting effects which will be too small to affect the dynamic behaviour. In the second part of the paper we shall set out the calculation process and then demonstrate the validity of our model by comparing it with experimental results.

The geometric model gave some interesting insights into the motion of the piston within the chamber. The pressures across the piston, resulting in forces and movements provided useful insights for the detailed modelling.

The identification of the effect of pipeline length and the resulting variation in flow rate had not, to our knowledge, been reported in the literature before in the context of this type of meter. Another effect of the flow, which proves to be a significant contribution to the pressure loss across the meter, is the loss at entry and exit. This is, also, a difficult loss to model, and the approach we took made use of losses measured for pipework features which approximated those in the meter. However, it proved necessary to make use of experimental data also.

Another force causing deceleration of the piston was found to be the trapped liquid at TDC and BDC. Clearly this was known to manufacturers who had introduced various cut-outs to release this trapped liquid. Again, the modelling of the flow losses required some estimation and would be an area for computational fluid dynamic modelling if the improvement was considered worthwhile.

The effect of lubrication holes on the lift of the piston will be discussed in Part II of this paper. For the calculation of the resultant tilt of the piston from squeeze films we made use of an approximate solution [16-18]. Further work might address this although, as we shall show, the predictions agree well with the data.

We consider that our friction model was well founded and provided the final element of the modelling.

For details of the leakage model the reader is referred to reference [11].

The numerical model is a time marching solution which applies the forces at each position and deduces the new position and velocities. Thus the local size of the various forces is calculated at each point in the motion for the precise position of the piston at that moment.

8. Conclusions

In this paper we have derived in detail the forces which are necessary to allow a solution of the equations of motion which make up the numerical model. These forces include:

- external pressures;

- internal pressures;

- friction forces;

- trapped liquid in the mechanism;

- pressure losses in the pipework;

- unsteady effects.

Part II of this paper will demonstrate the success of the model, which has now been applied to a smaller design of meter for production with success.

Acknowledgements

CEM was supported by the UK Engineering and Physical Sciences Research Council (EPSRC) and Litre Meter through an Industrial CASE Studentship. RCB received support from the Gatsby Charitable Foundation.

# References

- Baker, P. D. (1983) "Positive Displacement Liquid Meters." Technical Paper 101B Smith Meter Inc: 1-6.

- Ambrosius, E. E. (1948). "Analysis of tests on positive displacement meters for liquid hydrocarbons." Transactions of the ASME November: 703-708.

- Hahn, B. V. (1968). "Theory of the sliding vane meter." Siemens Rev. XXXV(9): 362-6.

- Baker, R. C. and Morris, M. V. (1985) "Positive displacement meters for liquids." Trans Inst M C 7(4): 209-220.

- Morton, C. E., Baker, R. C. and Hutchings, I. M. (2011) Measurement of liquid film thickness by optical fluorescence and its application to an oscillating piston positive displacement flowmeter. Measurement Science & Technology. 22 125403 (11pp).

- Cloran, E. M. (1938). "Rotary Oil Meters of the Positive Displacement and Current Types." Trans ASME: 617-624.

- Parry, A. (2001). "Computation of fluid-rigid body dynamic interaction in flowmeters." Flow Measurement 2001-International Conference, National Engineering Laboratory.

- Morton, C. E., Hutchings, I. M. and Baker, R. C. (2014a) Experimental investigation of an oscillating circular piston positive displacement flowmeter: I - piston movement and pressure losses. Flow Measurement and Instrumentation 36 (2014)47–56

- Morton, C. E., Baker, R. C. and Hutchings, I. M. (2014b) Experimental investigation of an oscillating circular piston positive displacement flowmeter: II - leakage flows and wear tests. Flow Measurement and Instrumentation 36 (2014) 57–63

- Pars, L. A. (1953). Introduction to Dynamics, Cambridge University Press.

- Morton, C. E. (2009) Performance and modelling of the oscillating piston flowmeter. PhD Thesis, Engineering Department, University of Cambridge.

- Hamrock, B. J. (1994). Fundamentals of fluid film lubrication, Mc Graw Hill.

- Chadwick, A., J. Morfett, et al. (2004). Hydraulics in civil and environmental engineering, Spoon Press.

- Miller, D. S. (1990). Internal flow systems; Design and Performance Predictions, Miller Innovations.

- Cameron, A. (1981). Basic Lubrication Theory, Ellis Horwood.

- Etsion, I. (1980). "Squeeze Effects in Radial Face Seals." Journal of Lubrication Technology 102: 145-151.

- Etsion, I. (1978). "Nonaxisymmetric Incompressible Hydrostatic Pressure Effects in Radial Face Seals." Journal of Lubrication Technology 100: 379-385.

- Etsion, I. (1979). "Hydrodynamic Effects in a Misaligned Radial Face Seal." Journal of Lubrication Technology 101: 283-293.