- Original Research Article -

Effect of a calcium deblooming algorithm on accuracy of coronary computed tomography angiography.

Jonathan R Weir-McCall1, Rui Wang2, Jaydeep Halankar3, Jiang Hsieh4, Cameron J Hague3, Samuel Rosenblatt3, Zhanming Fan2, Stephanie Sellers3, Darra T Murphy3, Philipp Blanke3, Lei Xu2*, Jonathon A Leipsic3

1.Department of Radiology, University of Cambridge, Cambridge

2. Department of Radiology, Beijing Anzhen Hospital, Capital Medical University, Beijing

3. St Pauls Hospital, Vancouver

4. GE Healthcare Technologies, Waukesha, Wisconsin

Short title: Effect of calcium deblooming on CTA accuracy

Funding: This study was supported by grant from National Key R&D Program of China (2016YFC1300300) and grant from Natural Science Foundation of China (81641069).

Disclosures: Dr. Leipsic serves as a consultant and has stock options in HeartFlow and Circle Cardiovascular Imaging, and receives speaking fees from GE Healthcare. All other authors have no conflicts of interest to disclose.

Word Count: 2347

Corresponding Author*:

Lei Xu, M.D.,PH.D

Department of Radiology, Beijing Anzhen Hospital, Capital Medical University.

No 2, Anzhen Road, Chaoyang District, Beijing 100029, China

Email: leixu2001@hotmail.com

Tel: 86-10-64456071

TOC Summary: Coronary calcification deblooming using a prototype post-processing algorithm is feasible and significantly reduces calcium blooming with an improvement of the specificity of the CTA exam.

Abstract

Background: Coronary artery calcification is a significant contributor to reduced accuracy of coronary computed tomographic angiography (CTA) in the assessment of coronary artery disease severity.  The aim of the current study is to assess the impact of a prototype calcium deblooming algorithm on the diagnostic accuracy of CTA.

Methods: 40 patients referred for invasive catheter angiography underwent CTA and invasive catheter angiography. The CTA were reconstructed using a standard soft tissue kernel (CTASTAND) and a deblooming algorithm (CTADEBLOOM).  CTA studies were read with and without the deblooming algorithm blinded to the invasive coronary angiogram findings.  Sensitivity, specificity, accuracy, positive predictive value and negative predictive value for the detection of stenosis ≥50% or ≥70% were evaluated using quantitative coronary angiography as the reference standard.  Image quality was assessed using a 5-point scale, and the presence of image artifact recorded.

Results: All studies were diagnostic with 548 segments available for evaluation.  Image score was 3.64±0.72 with CTADEBLOOM, versus 3.56±0.72 with CTASTAND (p=0.38). CTADEBLOOM had significantly less calcium blooming artifact than CTASTAND (12.5% vs. 47.5%, p=0.001).  Based on a 50% stenosis threshold for defining significant disease, the Sensitivity/Specificity/PPV/NPV/Accuracy were 65.9/84.9/27.6/96.6/83.4 for CTADEBLOOM and 75.0/81.9/26.6/97.4/81.4 for CTASTAND using a ≥50% threshold.  CTADEBLOOM specificity was significantly higher than CTASTAND (84.9% vs. 81.5%, p=0.03), with no difference between the algorithms in sensitivity (p=0.22), or accuracy (p= 0.15). These results remained unchanged when a stenosis threshold of ≥70% was used. Interobserver agreement was fair with both techniques (CTADEBLOOM k= 0.38, CTASTAND k= 0.37).

Conclusion: In this proof of concept study, coronary calcification deblooming using a prototype post-processing algorithm is feasible and reduces calcium blooming with an improvement of the specificity of the CTA exam.

Introduction

Coronary computed tomographic angiography (CTA) is currently indicated for the investigation of stable chest pain in low-intermediate risk patients in the American College of Cardiology and European Society of Cardiology guidelines.1,2  This is due to the high sensitivity and negative predictive value evidenced in multiple diagnostic accuracy studies, allowing for the effective exclusion of coronary artery disease following a negative test.3–5  Indeed a recent meta-analysis has shown CT to be the best test to rule out coronary artery disease (CAD) irrespective of pre-test probability.6

The ability of CT to rule out significant CAD is significantly impacted by the presence of coronary artery calcification.  Even moderate levels of coronary calcification are associated with a reduction in the ability to exclude significant stenosis.7 The challenge of dealing with calcification on CTA is only going to increase in the coming years due to an aging population and the known association of age with coronary calcification.8 Techniques to reduce the impact of coronary calcification on the accuracy of CTA are thus required.

Calcium reduces the accuracy of coronary assessment secondary to partial volume artifact amplifying the apparent size of the calcific plaque – also referred to as calcium ‘blooming’.  A recent study has demonstrated the potential to reduce the effect of calcium blooming through the use of a post-processing algorithm9, however the effect of this on the accuracy of CTA has not been assessed. The aim of this current proof of concept study is to test the impact of this prototype deblooming algorithm on the accuracy of CTA for the detection of significant coronary artery stenosis compared with the reference standard of invasive quantitative coronary angiography (QCA).  The study hypothesis was that calcium deblooming would increase diagnostic accuracy due to improved stenosis assessment in a cohort of participants with significant calcium burden.

Methods

Study design

Population

This single centre retrospective study recruited consecutive patients who were referred for ICA for the assessment of known or suspected CAD who had CTA prior to their ICA.  Patients with a coronary artery calcium score of greater than 100 were included. Patients with no calcified plaque, or with previous coronary artery bypass grafting or coronary stenting were excluded from this study group.

Ethics approval was granted by the Beijing Anzhen Hospital Ethics Committee. Informed consent was waived by the institutional review board due to the purely retrospective nature of the study. All study procedures were conducted in accordance with the Declaration of Helsinki.

CTA acquisition and postprocessing

All CTAs were performed at a single institution using a 256-row detector CT scanner (Revolution CT, GE healthcare, Waukasha, WI) with prospectively ECG-triggering volume scan within a single cardiac cycle. Tube voltage and current were automatically selected by kV-Assist and by Smart-mA based on the scout image. The data acquisition window was set according to HR. Contrast medium of 60–65 ml (370 mg iodine/ml, iopromide; Ultravist, Bayer Schering Pharma, Berlin, Germany) was injected at 4–5 ml/s rate followed by 30 ml of saline with a dual-head power injector. Scanning parameters were as follows: 256 × 0.625 mm collimation, and scan coverage was 120 or 140 mm with a matrix size of 512 × 512 pixels and reconstruction section thickness and section interval of 0.625 mm.  Images were reconstructed using iterative reconstruction at 50% strength (ASIR-V) with a standard soft tissue convolution kernel. The DICOM of the best image phase (CTASTAND) was then used with a proprietary deblooming algorithm (CTADEBLOOM)9. The algorithm is based on the observation that the blooming effect is caused by the spread of high-density structures to the surrounding regions as a result of the system point spread function and beam hardening.  The algorithm first identifies the affected regions by segmenting calcified plaques, stents, and other high-density objects.  An iterative correction approach is then employed to restore the underlying signal affected by these objects. The algorithm is iterative in nature in order to overcome many shortcomings of the traditional deconvolution approaches.10

CTA Image Analysis

Image analysis was performed in an independent corelab.  The corelab was blinded to all clinical details, original CTA report, and ICA data.  All scans were reviewed by two experienced cardiac imagers (JL with 15 years of experience in CTA, and JWM with 7 years of experience in CTA), who were also blinded to the reconstruction algorithm.  All datasets were reviewed for image quality, the presence of artifacts and segmental stenosis analysis independently by both reviewers.

Global image quality was scored using a 5-point Likert scale (1 – Non-diagnostic image quality with significant artifacts, 2 – Poor image quality with moderate artifact but still diagnostic, 3 – Fair image quality with minor artifacts, 4 – good image quality with minimal artifacts, 5- excellent image quality free of artifacts).11  The presence or absence of the following artifacts was qualitatively assessed and recorded: Motion, Calcium, Low intravascular contrast, and Noise.11 Image quality score was averaged between the two reviewers.  If an artifact was identified by either of the two reviewers it was considered to be present.

The CTAs were analysed using dedicated software on a single workstation (ADW 4.5, GE, Piscataway, New Jersey). A segmental analysis was performed and recorded using an 18 segment coronary arterial model.12  Each segment was scored as having significant coronary artery disease if there was a stenosis ≥50%.  Absent segments were not scored.  Segments which were not assessable were scored as having significant coronary artery disease. Where there was a disagreement between the two readers, a third reader (CJH – 10 years experience in CTA reporting) adjudicated on the presence/absence of significant coronary artery disease.

ICA performance

All scans were carried out within one month of CTA.  All ICA were performed by 2 practitioners.

ICA analysis

One experienced cardiovascular imager (RW, with 10 years of experience in cardiac imaging), blinded to the results of the CTA corelab, identified and analyzed all coronary segments, using the same 18-segment model as used for the CTA reporting. All segments were visually inspected for the presence of stenosis in at least 2 orthogonal views.  Where a ≥30% diameter stenosis was present on at least one view, QCA was performed using commercially available post-processing software (Medis 2.0). Where a total occlusion was present, only the most proximal segment was scored, with those distally not scored.  Presence or absence of stenosis ≥50% on QCA was used as the reference standard for comparison with CTA.13

Statistical analysis

Continuous variables are expressed as means ± standard deviation (SD), and categorical variables are presented as absolute counts and percentages. Between group differences were analyzed using the McNemar’s test for categorical variables, and those between continuous variables using a paired Student’s t-test. Sensitivity, specificity, accuracy, positive predictive value and negative predictive values were calculated for CTA using a standard post processing algorithm and a deblooming algorithm on a per patient, per vessel and per segment level using QCA as the reference standard.  Differences in sensitivity, specificity and accuracy between the algorithms were compared using McNemar’s test.  A sensitivity analysis was performed using a threshold of a diameter stenosis of 70% on ICA and CTA as being considered significant.  Cohen kappa was used to assess inter-observer variability, with a Cohen’s κ statistic of ≤0.2 = poor agreement, 0.21–0.4 = fair agreement, 0.41–0.6 = moderate agreement, 0.61–0.8 = good agreement and 0.81–1 = excellent agreement. All analyses were performed with SPSS (version 22, IBM SPSS, NY).

Results:

Study population

40 patients (mean age 61±7 y, 9 (22.5%) female) met the study inclusion criteria.  The demographic data of the study cohort is summarized in Table 1.  55% demonstrated no evidence of obstructive coronary artery disease, whilst 30% had single vessel disease, 10% had two vessel disease, and 5% had left main disease/three vessel disease on ICA.  No adverse events were experienced during the study using either diagnostic test. The STARD flow diagram of the study participants is shown in Figure 1.

CTA Image quality

The mean image quality score was 3.64±0.72 with CTADEBLOOM, versus 3.56±0.72 with CTASTAND (p=0.38). There were no significant differences in the presence of motion artifact (CTADEBLOOM 47.5% vs. 55% CTASTAND, p=0.51) or image noise (CTADEBLOOM 32.5% vs. 25% CTASTAND, p=0.58).  CTADEBLOOM suffered from significantly less calcium blooming artifact (CTADEBLOOM 12.5% vs. 47.5% CTASTAND, p=0.001, Figure 2), but demonstrated a significantly higher rate of intraluminal artifact related to the post-processing technique (CTADEBLOOM 17.5% vs. 0% CTASTAND, p=0.016, see Figure 3).

Per Segment analysis

Of 720 theoretically possible segments (based on the 18 segment model and 40 patients) 150 segments were not present, leaving 570 scored segments on CTA analysis.  Of the 570, 20 were scored as obstructive due to being non-evaluable on CTADEBLOOM, and 18 on CTASTAND. 22 segments were excluded from the final analysis as they lay distal to a complete obstruction on ICA.  This left 548 segments in the final comparative analysis.

The Sensitivity/Specificity/Accuracy/PPV/NPV were 65.9/84.9/83.4/27.6/96.6 for CTADEBLOOM and 75.0/81.9/81.4/26.6/97.4 for CTASTAND (Table 2).  The CTADEBLOOM algorithm was significantly more specific than CTASTAND for the detection of significant coronary artery disease (CTADEBLOOM 84.9% vs. CTASTAND 81.5%, p=0.03), with no significant difference between the algorithms in sensitivity (p=0.22), or accuracy (p= 0.15).

There was no significant change in these observations when a diameter stenosis threshold of ≥70% was considered obstructive.  At this threshold, the Sensitivity/Specificity/Accuracy/PPV/NPV were 61.5/92.0/90.5/27.6/98.0 for CTADEBLOOM and 69.2/89.3/88.4/24.3/98.4 for CTASTAND.  The CTADEBLOOM algorithm was significantly more specific than CTASTAND for the detection of significant coronary artery disease (CTADEBLOOM 92.0% vs. CTASTAND 89.3%, p=0.03), with no significant difference between the algorithms in sensitivity (p=0.63), or accuracy (p= 0.07).

Per Vessel analysis

120 vessels were included in the analysis with none excluded.  The Sensitivity/Specificity/PPV/NPV were 86.8/64.6/53.2/91.4 for CTADEBLOOM and 92.1/59.8/51.5/94.2 for CTASTAND (Table 2). At a vessel level analysis there was no statistically significant difference between the algorithms in sensitivity (p=0.50), specificity (p=0.42) or overall accuracy (p= 0.80).

Per Patient analysis

All 40 participants were included in the analysis with none excluded.  The Sensitivity/Specificity/PPV/NPV were 88.9/23.1/70.6/50 for CTADEBLOOM and 100/23.1/73.0/100 for CTASTAND (Table 2). At a participant level analysis there was no statistically significant difference between the algorithms in sensitivity (p=0.25), specificity (p=1.0) or overall accuracy (p= 0.25).

Agreement

Interobserver per segment agreement was fair using both CTADEBLOOM (Kappa = 0.38 (SE = 0.04)), and for CTASTAND (Kappa = 0.37 (SE = 0.04)).

Discussion

This study provides proof of concept of the feasibility of the deblooming algorithm, with use of the algorithm resulting in a significant reduction in the presence of calcium blooming artifact, with a resultant improvement in the specificity of the CTA analysis. These results remained robust to whether a stenosis of 50% or 70% was considered obstructive.

Multiple studies have documented the negative impact of coronary artery calcium scoring on the accuracy of CTA.14–18 Coronary artery calcification is impactful in two manners: first it reduces the ability to accurately visualize the lumen and second the presence of severe coronary artery calcification is a marker of an increased pretest probability of having a significant stenosis.19 Due to the deleterious effect of calcification on diagnostic yield it has been proposed that above a set calcium threshold (typically 600-1000 Agatston Units), CTA is not performed and the patient is referred directly onwards for further evaluation.  The efficacy of such an approach has been demonstrated in a randomized clinical trial20, yet even in those with a calcium score >1200, CTA is able to effectively exclude obstructive coronary artery disease in 37% of patients.7 With an aging population and an increasing prevalence of primary prevention with statin therapy - both of which are known to cause a rise in coronary artery calcification - the need to better cope with calcification is necessary.21

Several approaches have been assessed in CTA utilizing advances in hardware, and advanced post-processing techniques.  By increasing the spatial resolution of CT the impact of blooming is reduced due to the reduction in partial voluming.  A study by Pontone et al. demonstrated that reducing the x-y axis resolution from 0.625mm to 0.23mm improved the accuracy of calcified plaque evaluation.22  An improvement in the x-y axis resolution however does not reduce calcium blooming in the z-axis which is dependent on scanner slice thickness which has not substantially changed in the last 10 years.  A clinical prototype scanner has recently been developed with an improved z-axis resolution of 0.23mm which may overcome this limitation.  Preliminary reports describe good accuracy compared with ICA, with successful reduction in calcium blooming and a reduction of false positives from 8 to 1 in 24 calcified segments using high resolution images compared with standard resolution images.23,24 Calcium subtraction has been useful in peripheral CT angiography, but has been limited in CTA due to the need to acquire two scans of the heart resulting in both high radiation doses, and misregistration artifacts.25 Using a dual source CT scanner, De Santis and colleagues have recently been able to use dual energy CT to good effect, with improved lumen visualization and image quality reported compared with the unsubtracted images.26  However the impact of this on diagnostic accuracy has yet to be evaluated.

High-convolution filter reconstruction algorithms have been used with good efficacy in stent evaluation, and hold promise in calcium blooming reduction although are associated with a significant increase in image noise which is problematic in those with an elevated BMI.27,28 This may in future be overcome through the combination of high convolution filters with model based iterative reconstruction which improves the signal to noise ratio as well as reducing the volume of calcification.29,30  Direct comparison of these techniques with the current deblooming algorithm will be useful in future iterations of the deblooming algorithm.

In the current study we examined a prototype calcium deblooming algorithm.  An advantage of this over hardware advances is its lack of requirement for new equipment, and can therefore be rapidly adopted and retrofitted to already existent scanners.  Unlike high convolution filter techniques which increase image noise, it has previously been shown that the current technique both reduces calcium volume and increases the signal to noise ratio.9  In the current study we found no difference between the two techniques in the presence of image noise.  There was a significant reduction in the subjective degree of calcium blooming, but at the cost of the occurrence of an intraluminal artifact in 17.5% of the studies.  This intra-luminal artifact may account for the fact that while a significant improvement in the specificity with the deblooming algorithm, this was not accompanied by a significant improvement in the overall diagnostic accuracy.  Overall the inter-observer agreement on the presence of significant stenosis was only fair using both techniques with Kappa values of 0.37-0.38.  This is comparable with those previous reported by Williams et al., in the SCOT-HEART cohort with kappa values of 0.46 on per segmental analysis.31 This highlights a need to continue to improve on our current visual analysis techniques.

There are several limitations to the current study.  It is a single centre study and as such it is limited by all the inherent biases that are associated with this.  The cohort studied were those scheduled for ICA which could lead to an inclusion bias.  Rate control was not used, nor was motion correction algorithms, as a result of which a significant number of scans suffered from motion artifact.  Despite this, only 20 out of 581 segments were non-evaluable showing that with current generation scanners that high accuracy is possible even without aggressive rate control.32 Further work is required to determine if the combination of the deblooming algorithm and motion correction results in further improvements in accuracy.  Due to the recruitment criteria there were very few stenotic plaques without calcification within them, thus a comparison of calcified versus non-calcified plaque could not be performed. In future larger studies it would be useful to examine this as well as to combine this with plaque quantification to examine the effect of increasing plaque calcification on lumen visibility. Finally, we used a 50% threshold for the definition of significant coronary artery disease.  Traditionally a 70% diameter stenosis on invasive coronary angiography has been defined as being obstructive. However in the post-FAME era, such an anatomical cut point is less relevant, with significant numbers of those with a 50-70% stenosis being flow limiting, and a smaller but equally significant fraction of those with >70% stenosis being non-flow limiting.33,34  The more clinically important questions for CTA are thus to first, identify or exclude coronary atherosclerosis, and second determine if further functional or invasive assessment is required, for which the 50% stenosis threshold is more appropriate.35,36

Conclusion

In this proof of concept study coronary calcification deblooming using a post-processing algorithm is feasible and significantly reduces calcium blooming with an improvement of the specificity of the CTA exam.

References

1. 	Fihn SD, Gardin JM, Abrams J, et al. 2012 ACCF/AHA/ACP/AATS/PCNA/SCAI/STS guideline for the diagnosis and management of patients with stable ischemic heart disease. J Am Coll Cardiol. 2012;60(24):e44-e164. doi:10.1016/j.jacc.2012.07.013.

2. 	Montalescot G, Sechtem U, Achenbach S, et al. 2013 ESC guidelines on the management of stable coronary artery disease. Eur Heart J. 2013;34(38):2949-3003. doi:10.1093/eurheartj/eht296.

3. 	Budoff MJ, Li D, Kazerooni EA, Thomas GS, Mieres JH, Shaw LJ. Diagnostic Accuracy of Noninvasive 64-row Computed Tomographic Coronary Angiography (CCTA) Compared with Myocardial Perfusion Imaging (MPI): The PICTURE Study, A Prospective Multicenter Trial. Acad Radiol. 2017;24(1):22-29. doi:10.1016/j.acra.2016.09.008.

4. 	Neglia D, Rovai D, Caselli C, et al. Detection of significant coronary artery disease by noninvasive anatomical and functional imaging. Circ Cardiovasc Imaging. 2015;8(3). doi:10.1161/CIRCIMAGING.114.002179.

5. 	Schuetz GM, Schlattmann P, Dewey M. Use of 3x2 tables with an intention to diagnose approach to assess clinical performance of diagnostic tests: meta-analytical evaluation of coronary CT angiography studies. BMJ. 2012;345(7881):e6717. doi:10.1136/bmj.e6717.

6. 	Knuuti J, Ballo H, Juarez-Orozco LE, et al. The performance of non-invasive tests to rule-in and rule-out significant coronary artery stenosis in patients with stable angina: a meta-analysis focused on post-test disease probability. Eur Heart J. 2018;39(35):3322-3330. doi:10.1093/eurheartj/ehy267.

7. 	Schuhbaeck A, Schmid J, Zimmer T, et al. Influence of the coronary calcium score on the ability to rule out coronary artery stenoses by coronary CT angiography in patients with suspected coronary artery disease. J Cardiovasc Comput Tomogr. 2016;10(5):343-350. doi:10.1016/j.jcct.2016.07.014.

8. 	McClelland RL, Chung H, Detrano R, Post W, Kronmal RA. Distribution of coronary artery calcium by race, gender, and age: Results from the Multi-Ethnic Study of Atherosclerosis (MESA). Circulation. 2006;113(1):30-37. doi:10.1161/CIRCULATIONAHA.105.580696.

9. 	Li P, Xu L, Yang L, et al. Blooming Artifact Reduction in Coronary Artery Calcification by A New De-blooming Algorithm: Initial Study. Sci Rep. 2018;8(1):1-8. doi:10.1038/s41598-018-25352-5.

10. 	Hsieh J, Dutta S. System and method for quality improvement in CT imaging, United States Patent: 8699812. 2011. http://patft.uspto.gov/netacgi/nph-Parser?Sect1=PTO1&Sect2=HITOFF&d=PALL&p=1&u=%2Fnetahtml%2FPTO%2Fsrchnum.htm&r=1&f=G&l=50&s1=8,699,812.PN.&OS=PN/8,699,812&RS=PN/8,699,812. Accessed April 12, 2019.

11. 	Nakanishi R, Sankaran S, Grady L, et al. Automated estimation of image quality for coronary computed tomographic angiography using machine learning. Eur Radiol. 2018;28(9):4018-4026. doi:10.1007/s00330-018-5348-8.

12. 	Leipsic J, Abbara S, Achenbach S, et al. SCCT guidelines for the interpretation and reporting of coronary CT angiography: a report of the Society of Cardiovascular Computed Tomography Guidelines Committee. J Cardiovasc Comput Tomogr. 2015;8(5):342-358. doi:10.1016/j.jcct.2014.07.003.

13. 	Meijboom WB, Meijs MFL, Schuijf JD, et al. Diagnostic Accuracy of 64-Slice Computed Tomography Coronary Angiography. A Prospective, Multicenter, Multivendor Study. J Am Coll Cardiol. 2008;52(25):2135-2144. doi:10.1016/j.jacc.2008.08.058.

14. 	Dewey M, Clouse ME, Bush DE, Hoe JWM, Miller JM. Coronary Artery Stenoses : Accuracy of 64 – Detector Row CT Angiography in Segments with Mild , Subanalysis of the CORE-64 Trial 1 Methods : Results : Radiology. 2011;261(1):100-108. doi:10.1148/radiol.11110537/-/DC1.

15. 	Ong TK, Chin SP, Liew CK, et al. Accuracy of 64-row multidetector computed tomography in detecting coronary artery disease in 134 symptomatic patients: Influence of calcification. Am Heart J. 2006;151(6):4-9. doi:10.1016/j.ahj.2005.12.027.

16. 	Palumbo AA, Maffei E, Martini C, et al. Coronary calcium score as gatekeeper for 64-slice computed tomography coronary angiography in patients with chest pain: Per-segment and per-patient analysis. Eur Radiol. 2009;19(9):2127-2135. doi:10.1007/s00330-009-1398-2.

17. 	Ahn SJ, Kang DK, Sun JS, Yoon MH. Accuracy and predictive value of coronary computed tomography angiography for the detection of obstructive coronary heart disease in patients with an agatston calcium score above 400. J Comput Assist Tomogr. 2013;37(3):387-394. doi:http://dx.doi.org/10.1097/RCT.0b013e318282d61c.

18. 	Qi L, Tang L-J, Xu Y, et al. The Diagnostic Performance of Coronary CT Angiography for the Assessment of Coronary Stenosis in Calcified Plaque. PLoS One. 2016;11(5):e0154852. doi:10.1371/journal.pone.0154852.

19. 	Mitsutake R, Niimura H, Miura S, et al. Clinical significance of the coronary calcification score by multidetector row computed tomography for the evaluation of coronary stenosis in Japanese patients. Circ J. 2006;70(9):1122-1127. http://www.ncbi.nlm.nih.gov/pubmed/16936422.

20. 	Lubbers M, Dedic A, Coenen A, et al. Calcium imaging and selective computed tomography angiography in comparison to functional testing for suspected coronary artery disease: The multicentre, randomized CRESCENT trial. Eur Heart J. 2016;37(15):1232-1243. doi:10.1093/eurheartj/ehv700.

21. 	Lee S-E, Chang H-J, Sung JM, et al. Effects of Statins on Coronary Atherosclerotic Plaques: The PARADIGM (Progression of AtheRosclerotic PlAque DetermIned by Computed TomoGraphic Angiography Imaging) Study. JACC Cardiovasc Imaging. June 2018. doi:10.1016/j.jcmg.2018.04.015.

22. 	Pontone G, Bertella E, Mushtaq S, et al. Coronary Artery Disease: Diagnostic Accuracy of CT Coronary Angiography—A Comparison of High and Standard Spatial Resolution Scanning. Radiology. 2014;271(3):688-694. doi:10.1148/radiol.13130909.

23. 	Motoyama S, Ito H, Sarai M, et al. Ultra-High-Resolution Computed Tomography Angiography for Assessment of Coronary Artery Stenosis. Circ J. 2018;82(7):1844-1851. doi:10.1253/circj.CJ-17-1281.

24. 	Takagi H, Tanaka R, Nagata K, et al. Diagnostic performance of coronary CT angiography with ultra-high-resolution CT: Comparison with invasive coronary angiography. Eur J Radiol. 2018;101(January):30-37. doi:10.1016/j.ejrad.2018.01.030.

25. 	Yoshioka K, Tanaka R, Muranaka K, et al. Subtraction coronary CT angiography using second-generation 320-detector row CT. Int J Cardiovasc Imaging. 2015;31(1):51-58. doi:10.1007/s10554-015-0630-1.

26. 	De Santis D, Jin KN, Joseph Schoepf U, et al. Heavily calcified coronary arteries advanced calcium subtraction improves luminal visualization and diagnostic confidence in dual-energy coronary computed tomography angiography. Invest Radiol. 2018;53(2):103-109. doi:10.1097/RLI.0000000000000416.

27. 	Cui X, Li T, Li X, Zhou W. High-definition computed tomography for coronary artery stents imaging: Initial evaluation of the optimal reconstruction algorithm. Eur J Radiol. 2015;84(5):834-839. doi:10.1016/j.ejrad.2015.02.004.

28. 	Sun Z, Ng CKC, Xu L, Fan Z, Lei J. Coronary ct angiography in heavily calcified coronary arteries: Improvement of coronary lumen visualization and coronary stenosis assessment with image postprocessing methods. Med (United States). 2015;94(48):1-10. doi:10.1097/MD.0000000000002148.

29. 	Precht H, Kitslaar PH, Broersen A, et al. Influence of Adaptive Statistical Iterative Reconstruction on coronary plaque analysis in coronary computed tomography angiography. J Cardiovasc Comput Tomogr. 2016;10(6):507-516. doi:10.1016/J.JCCT.2016.09.006.

30. 	Károlyi M, Szilveszter B, Kolossváry M, et al. Iterative model reconstruction reduces calcified plaque volume in coronary CT angiography. Eur J Radiol. 2017;87:83-89. doi:10.1016/j.ejrad.2016.12.012.

31. 	Williams MC, Golay SK, Hunter  a., et al. Observer variability in the assessment of CT coronary angiography and coronary artery calcium score: substudy of the Scottish COmputed Tomography of the HEART (SCOT-HEART) trial. Open Hear. 2015;2(1):e000234-e000234. doi:10.1136/openhrt-2014-000234.

32. 	Liang J, Wang H, Xu L, et al. Diagnostic performance of 256-row detector coronary CT angiography in patients with high heart rates within a single cardiac cycle: a preliminary study. Clin Radiol. 2017;72(8):694.e7-694.e14. doi:10.1016/j.crad.2017.03.004.

33. 	Pijls NHJ, de Bruyne B, Peels K, et al. Measurement of Fractional Flow Reserve to Assess the Functional Severity of Coronary-Artery Stenoses. N Engl J Med. 1996;334(26):1703-1708. doi:10.1056/NEJM199606273342604.

34. 	Curzen N, Rana O, Nicholas Z, et al. Does routine pressure wire assessment influence management strategy at coronary angiography for diagnosis of chest pain? the ripcord study. Circ Cardiovasc Interv. 2014;7(2):248-255. doi:10.1161/CIRCINTERVENTIONS.113.000978.

35. 	Curzen NP, Nolan J, Zaman AG, Nørgaard BL, Rajani R. Does the Routine Availability of CT–Derived FFR Influence Management of Patients With Stable Chest Pain Compared to CT Angiography Alone?: The FFRCT RIPCORD Study. JACC Cardiovasc Imaging. 2016;9(10):1188-1194. doi:10.1016/j.jcmg.2015.12.026.

36. 	Cury RC, Abbara S, Achenbach S, et al. CAD-RADS(TM) Coronary Artery Disease - Reporting and Data System. An expert consensus document of the Society of Cardiovascular Computed Tomography (SCCT), the American College of Radiology (ACR) and the North American Society for Cardiovascular Imaging (. J Cardiovasc Comput Tomogr. 2016;10(4):269-281. doi:10.1016/j.jcct.2016.04.005.

Figure 1: STARD participant flow diagram for the study

Figure 2: Comparison of the left anterior descending artery using a standard (left) and deblooming (right) reconstruction algorithm.  Compared with the standard reconstruction on the left there is a significant reduction in the degree of blooming with a resultant reduction in the severity of the stenosis proximally on the right.

Figure 3: Example of the intra-luminal high attenuation artifact which was present in 17.5% of studies.

Table 1: Baseline patient characteristics

|  | Total |
| --- | --- |
| N | 40 |
| Age (years) | 61±7 |
| Sex (male) | 31m/9f |
| BMI (kg/m2) | 31±11 |
| SBP (mmHg) | 129±14 |
| DBP (mmHg) | 76±9 |
| DM | 14 |
| Hypertension | 27 |
| Hyperlipidaemia | 12 |
| Smoking Hx: |  |
| - Current | 13 |
| - Previous | 4 |
| - Never | 23 |
| FHx CVD | 15 |
| Obstructive disease on ICA* |  |
| No obstructive disease Single vessel Two vessel Left main/Three vessel | 22 (55%) 12 (30%) 4 (10%) 2 (5%) |

*LM≥50% or LAD/RCA/LCx with ≥70% diameter stenosis.

BMI – Body mass index; CAD – coronary artery disease; CACS – Coronary artery calcium score; DBP – Diastolic blood pressure; DM – Diabetes mellitus; FHx – Family history; ICA – Invasive Catheter Angiography; SBP – Systolic blood pressure.

Table 2:  Sensitivity, specificity, accuracy, positive predictive value and negative predictive value of CTA with (CTADEBLOOM) and without (CTASTAND) application of the calcification deblooming algorithm.

|  | TP | FP | TN | FN | Sensitivity | Specificity | Accuracy | Positive predictive value | Negative predictive value |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Per Segment |  |  |  |  |  |  |  |  |  |
| CTADEBLOOM | 29 | 76 | 428 | 15 | 65.9 | 84.9 | 83.4 | 27.6 | 96.6 |
| CTASTAND | 33 | 91 | 413 | 11 | 75 | 81.9 | 81.4 | 26.6 | 97.4 |
| Per Vessel |  |  |  |  |  |  |  |  |  |
| CTADEBLOOM | 33 | 29 | 53 | 5 | 86.8 | 64.6 | 71.7 | 53.2 | 91.4 |
| CTASTAND | 35 | 33 | 49 | 3 | 92.1 | 59.8 | 70.0 | 51.5 | 94.2 |
| Per Patient |  |  |  |  |  |  |  |  |  |
| CTADEBLOOM | 24 | 10 | 3 | 3 | 88.9 | 23.1 | 67.5 | 70.6 | 50.0 |
| CTASTAND | 27 | 10 | 3 | 0 | 100 | 23.1 | 75.0 | 73.0 | 100 |