Technology in the management of type 1 diabetes –

present status and future prospects

Martin Tauschmann1,2 and Roman Hovorka1,2

1Wellcome Trust-MRC Institute of Metabolic Science, University of Cambridge,

Cambridge, United Kingdom

2Department of Paediatrics, University of Cambridge, Cambridge, United Kingdom

Corresponding author:

Roman Hovorka, PhD
University of Cambridge Metabolic Research Laboratories
Level 4, Institute of Metabolic Science
Box 289, Addenbrooke’s Hospital, Hills Rd, Cambridge CB2 0QQ, UK
tel: +44 1223 762 862, fax:+44 1223 330 598,
email: rh347@cam.ac.uk

Word count: 				3988

Abstract:				179

Figures:				1

Tables:				4

References				130

Tables

- Table 1. Evidence supporting clinical use of pump therapy, continuous glucose monitoring, and glucose responsive insulin delivery in adults with type 1 diabetes

- Table 2. Evidence supporting clinical use of pump therapy, continuous glucose monitoring, and glucose responsive insulin delivery in children and adolescents with type 1 diabetes

- Table 3. ADA evidence-grading system for clinical practice recommendations

- Table 4. Professional society guidelines for  use of insulin pump therapy and continuous glucose monitoring in type 1 diabetes

Figures

- Figure 1. Currently available diabetes technologies in type 1 diabetes

Abstract

Type 1 diabetes represents 5-10% diabetes cases worldwide and is characterised by increasing incidence and no immediate prospect of cure. Lifelong management is required and eased by novel treatment modalities particularly from the field of diabetes technologies. Continuous glucose monitoring has become the standard of care including factory-calibrated subcutaneous glucose monitoring and long-term implantable glucose sensing.1 Significant progress has been made in technology-enabled glucose responsive insulin delivery. The first hybrid insulin-only closed-loop system has been commercialised and other closed-loop systems are under development including dual-hormone glucose control systems. The present review focuses on well-established diabetes technologies including glucose sensing, pen-based insulin delivery, data management and data analytics. We also cover insulin pump therapy, threshold-based suspend, predictive low glucose suspend, and single-hormone and dual-hormone closed-loop systems. Clinical practice recommendations for insulin pump therapy and continuous glucose monitoring are presented and ongoing research and future prospects highlighted. We conclude that the management of type 1 diabetes is enhanced by diabetes technology for the benefits of the majority of people with type 1 diabetes, their carers and guardians, and treating health care professionals.

179 words

.

Introduction

Type 1 diabetes represents 5-10% cases with diabetes worlwide.2 Despite intensive research, type 1 diabetes is presently incurable. Over the decades, diabetes management has greatly benefited from innovations in diabetes technology. The present era is witnessing an unprecedented scale of innovative diabetes technologies aimed to improve outcomes and ease the burden of diabetes management. Glucose monitoring has evolved from inaccurate bulky devices to factory-calibrated continuous glucose sensing connected to smartphones. Advances in insulin formulations and insulin delivery including insulin pump therapy and glucose responsive insulin delivery have led to effective insulin dosing. Software tools systematically track and manage complex glucose and insulin delivery data.

The present review covers established and novel diabetes technologies used in the management of type 1 diabetes. We describe presently available technologies and their impact on health outcomes including recommendations on their clinical use. We provide insights into diabetes technologies not yet widely adopted or under development.

The narrative is based on results from clinical studies and meta-analyses known to the authors. Relevant studies were identified by searching Pubmed and Web of Science for articles on each of the technologies up to November 2017, as well as using cited literature in retrieved articles. Information relevant to children, adolescents and adults with type 1 diabetes was considered; studies in pregnancy complicated by diabetes were excluded. Priority was given to meta-analyses and systematic reviews.

Widely adopted diabetes technologies

An overview of diabetes technologies and evidence supporting their use is provided in Table 1 and Table 2 using the level of evidence following the American Diabetes Association guidelines (Table 3).

Insulin delivery

Following the publication of the Diabetes Control and Complications Trial,3 the  treatment paradigm shifted towards intensive insulin therapy based on frequent blood glucose monitoring and flexible multiple daily insulin administrations through an insulin pen or an insulin pump.

Insulin pens

Insulin pens contain insulin in a cartridge and incorporate a fine replaceable needle. Introduced in 1981 as convenient, easy to use injection devices,4 pens are widely used as a part of multiple daily injection (MDI) therapy and are continuously evolving. Pens with memory functions (e.g. HumaPen® Memoir, Ely Lilly, Indianapolis, IN, USA; NovoPen Echo®, Novo Nordisk, Bagsværd, Denmark) or pen caps that track past doses (e.g. Timesulin®, Patients Pending Ltd, London, UK; GoCap®, Common Sensing, Cambridge, MA, USA) are available. Recently, pens with built-in Bluetooth connectivity have received regulatory approvals (e.g. InPen®, Companion Medical, San Diego, CA, USA; Esysta® pen, Emperra Digital Diabetes Care, Potsdam, Germany). These smart pens allow users to track doses, and automatically transfer data via Bluetooth to diabetes management apps residing on smartphones for convenience and automatic cloud upload for sharing data with health care professionals. However, no studies regarding superiority of smart pens over conventional pens have been reported.

Insulin pumps

Insulin pumps date back to 1970s5, but it took another 20 years for insulin pump therapy to become widely available. The increasing utilisation of insulin pump therapy over the last 20 years has resulted from improvements and increased reliability in pump technology, still yet imperfect, documented health benefits, and availability of rapid acting insulin analogues, and was further amplified by coverage by private insurance and public health care systems. Uptake and availability of insulin pump therapy varies considerably between and within countries6, with pump users representing 40 to 60% of the type 1 diabetes population in Western countries most adept to pump use.7, 8

Types of insulin pumps

Insulin pumps deliver short- or rapid acting insulin into the subcutaneous tissue at pre-programmed rates, normally half-hourly to hourly adjustable, with user-activated boosts/boluses at mealtimes via self-inserted Teflon or steel catheters. In conventional or tethered pumps, the pump’s insulin reservoir and the transcutaneously placed cannula are connected via tubing 18-42 inch long. Patch pumps comprise a very short insulin infusion set typically embedded inside the pump housing or within the base part of the modular designed pump9. While tethered pumps are usually tucked into pockets or carried in pump pouches, patch pumps are directly attached to the user’s skin. A recent retrospective observational study did not demonstrate any differences in HbA1c when comparing patch pumps vs. traditional tethered pumps.10

Adjunctive technologies

Modern insulin pumps usually come with adjunctive features such as bolus calculators to ease calculation of meal and correction boluses, bolus profiles including immediate and/or extended delivery of calculated bolus dose to meet postprandial insulin requirements, and temporary basal rates to accommodate physical activity resulting in acutely lower insulin needs, or stress or illness resulting in acutely higher insulin needs. Use of these advanced features may improve glycaemic outcomes including HbA1c11, and postprandial glycaemic excursions.12, 13

Efficacy of insulin pump therapy

In adults with type 1 diabetes, the use of insulin pump is associated with a modest 0.3 to 0.6% reduction in HbA1c compared to MDI therapy14-18, with those most poorly controlled on MDI experiencing the greatest, and often a substantial and clinically valuable improvement in HbA1c.14 A comparable to lower risk of severe hypoglycaemia has been documented, while quality of life in pump users is higher compared to MDI (Table 1)14-18. Despite the high appeal for children and adolescents due to more flexible and subtly customisable insulin delivery essential to paediatric needs, meta-analyses and systematic reviews of RCTs including paediatric populations14-17, 19 are not as conclusive as in adults. Similar to adults, slightly lower HbA1c levels and apparently no difference in severe hypoglycaemia risk were reported in meta-analyses of paediatric pump users compared to MDI therapy. Insulin requirements are usually lower on pump, while rates of DKA do not differ between pump and MDI therapy. Quality of life and treatment satisfaction is similar to higher than that of children and adolescents using MDI and their parents (Table 2).

With respect to severe hypoglycaemia, however, these meta-analyses should be interpreted with caution due to several issues such that clinical trials were too short for severe hypoglycaemia to occur or participation is limited intentionally or unintentionally to those with a very low rate of baseline hypoglycaemia. Severe hypoglycaemia or hypoglycaemia unawareness might have been listed as specific exclusion criteria in these trials, or early generation pumps and pump insulins were used. In a meta-regression analysis, Pickup et al. demonstrated that there is significant hypoglycaemia reduction even in children and adolescents on pump therapy compared to MDI, though this is to a lesser degree than in adults. Greatest reductions in severe hypoglycaemia occurred in those with highest baseline hypoglycaemia and in elderly.

While on the whole, the above mentioned meta-analyses of RCTs in paediatric and adult type 1 diabetes are cautiously favouring insulin pump therapy compared to MDI therapy, recent observational studies more optimistically documented sustained beneﬁt over prolonged periods of pump use across different populations including reductions in DKA and severe hypoglycaemia.20-25 This may reflect that non-minority and less social deprived users more frequently adopt insulin pump therapy.8, 26

Glucose monitoring

Capillary blood glucose measurements

Capillary measurement of blood glucose using handheld portable meters, glucose test strips and a lancet is most widely used. More frequent capillary testing correlates with improved HbA1c levels and reduced acute dysglycaemia.27, 28 Similar to bolus calculators residing on insulin pumps, recently introduced expert meters comprise integrated bolus advisors to calculate insulin dosages. Recent randomised controlled trials have shown a significant increase in the number of people achieving HbA1c targets29-31 in the bolus calculator group and a reduction in hypoglycaemia compared to controls.30, 31

Capillary blood glucose monitoring has its drawbacks as blood is sampled intermittently providing only snapshots of glucose concentrations even if performed frequently. Episodes of hyper- and hypoglycaemia may be missed and not factored into treatment decisions.

Continuous glucose monitoring

The emergence of continuous glucose monitoring (CGM) has been an important step for the glucose monitoring field. Currently available CGM devices measure interstitial glucose concentrations subcutaneously at one- to five-minute intervals utilizing enzyme-tipped electrodes or fluorescence technology. Readers, either stand-alone devices or integrated into insulin pumps or mobile phones, display transmitted interstitial glucose readings either in real-time (real-time CGM), or on demand when scanning (flash glucose monitoring), or simply collect data for retrospective read-out and analysis (professional, masked or blinded CGM).

Real-time CGM systems automatically display glucose readings at regular intervals, and utilise real-time alarms when sensor glucose levels reach pre-defined thresholds regarding hypo- and hyperglycaemia, as well as rate of change alarms for rapid glycaemic excursion. Recently introduced flash glucose monitoring systems (FreeStyle Libre, Abbott Diabetes Care, Alameda, CA, USA) report glucose levels only when the user scans the sensor by holding a reader or a cell phone close to the sensor. Blinded CGMs are applied intermittently over a short period of time to provide more information about glycaemic excursions and patterns to the healthcare professional in order to facilitate changes in therapy, and could serve as educational tools. Blinded CGM and flash glucose monitoring systems do not provide alarms.

While most CGMs still require calibration using capillary blood glucose readings, the Libre flash glucose monitoring system is factory calibrated and does not require re-calibration by the user.32 Most CGM systems are minimally invasive and have a life time of 6 to 14 days. A longer-term up to 6 months implantable sensor (Eversense®, Senseonics Inc, Germantown, MD, USA) is available in Europe.33 Sensor implantation and removal require a minor surgical procedure by a trained health care professional unlike for short term CGM systems which are self-inserted by the user.

CGM uptake and use

A niche product in the recent past, CGM has now become the standard of care for people with type 1 diabetes. Recent data from the DPV registry in Germany and Austria, and the T1D Exchange registry in the USA suggest that overall CGM use for all registry participants (DPV: n=20,938; T1D: n= 8,186) is 18.4% (DPV) and 21.7% (T1D), respectively34. Overall accuracy of the latest sensor generations measured as the mean relative absolute difference (MARD) versus a given laboratory standard is between 8% to14%32, 33, 35-37 with lower accuracy in the hypoglycaemic range and at rapidly changing glucose levels.38, 39 The technology has reached the proposed mark sufficient to allow self-adjustment of insulin dosage without confirmatory capillary blood glucose measurements (MARD <10%).40,41 CGM systems have received approvals for non-adjunctive use in the USA (G5™ Mobile, Dexcom, San Diego, California, USA; Libre Flash Glucose Monitor, Abbott Diabetes Care, Alameda, CA) and in the EU (G5™ Mobile, Dexcom; Libre Flash Glucose Monitor and FreeStyle Navigator II, Abbott Diabetes Care, Alameda, CA). Confirmatory capillary glucose measurement is suggested at hypoglycaemia with Libre.

Data provided by CGM devices allow limitations of the traditional glucose metrics such HbA1c and capillary glucose measurements to be overcome. A recent consensus report defined measures of glycaemic control based on CGM highlighting the significance of CGM technology in modern diabetes care.42, 43

Efficacy of CGM

RCTs and meta-analyses using early generation devices were cautious with respect to the overall benefit of CGM systems, particularly in children and young people with type 1 diabetes.15, 44-49 More recent data more consistently report that use of CGM is associated with an improvement in HbA1c, reduction in mild to moderate hypoglycaemia, and reduced glucose variability.50-55 While earlier analysis and guidelines were favouring CGM in combination with pump therapy15, 46, 56, 57, emerging evidence supports use of CGM as part of MDI.55, 58-60 As the technology is evolving fast, the older RTCs and meta-analyses have limited validity.

Flash glucose monitoring

With a 2-week sensor life, factory calibration, satisfactory accuracy with overall MARD of 11% to 14%, a small size and light weight, the recently introduced Libre flash glucose monitoring system is particularly appealing and convenient to assess glucose levels. 60, 61 However,  evidence on its effectiveness is limited.60, 62-64 An RCT showed that flash glucose monitoring in adults with well controlled type 1 diabetes reduced time spent in hypoglycaemia, reduced glucose variability, and improved time in range compared to self-monitoring of capillary blood glucose with median 15 scans per day.60 Benefits were identical for users of insulin pump therapy and MDI. In a recent head-to-head comparison of flash glucose monitoring and conventional CGM in adults with type 1 diabetes and impaired awareness of hypoglycaemia, CGM more effectively reduced time spent in hypoglycaemia compared with flash glucose monitoring. In the paediatric population, there is currently no evidence regarding effectiveness of flash glucose monitoring.61 Observational data link frequent scanning to improved outcomes.65

Glucose responsive insulin delivery

Automated suspension of insulin delivery at low glucose levels or when low glucose levels are predicted represent the early embodiments of technology-enabled glucose responsive regulation of insulin delivery to address the issue of hypoglycaemia. Closed-loop approaches are more complex and address both the issues of hypoglycaemia and hyperglycaemia.

Threshold-based insulin suspend

Released in 2009, the Medtronic Paradigm Veo (Medtronic Diabetes, Northridge, CA, USA) implements threshold-based insulin suspend. A revised version was approved in the USA in 2013 (MiniMed 530G).

Multicentre randomised controlled66-68 and non-randomised studies69-71 including in children and adolescents66, 67, 69 in real life settings have demonstrated that automated insulin suspension is safe and reduces the frequency and duration of overall and nocturnal hypoglycaemic episodes compared to insulin pump therapy alone66 or sensor augmented pump therapy67, 69 (Table 1 and Table 2) . Threshold-based suspend was shown to reduce the overall risk of severe and moderate hypoglycaemia in those with the highest risk, impaired hypoglycaemia awareness and the highest frequency of severe hypoglycaemia.66, 70

Predictive low glucose insulin suspend

Predictive low glucose insulin suspend discontinues insulin delivery when hypoglycaemia is predicted by an algorithm. This feature was introduced in Europe and Australia in 2015 (MiniMed 640G pump; Medtronic Diabetes). A revised version of this pump was approved in the USA for those aged sixteen and older (MiniMed 630G pump).

In RCTs including adults72, and children and adolescents72-74, the use of predictive low glucose suspend technology was shown to reduce the exposure to nocturnal72-74 and overall hypoglycaemia74, including reduced frequency of nocturnal and diurnal episodes and a reduction of prolonged nocturnal events. These benefits were achieved at the expense of mildly elevated overnight and morning glucose72, 73 or increased time in moderate hyperglycaemia.74

Closed-loop insulin delivery

Closed-loop systems (the artificial pancreas, or automated insulin delivery systems) are more elaborate and use a control algorithm that automatically and continually modulates insulin delivery below and above the pre-set rate based on sensor glucose levels.

Control algorithms used in academic and commercial closed-loop systems include a proportional-integral-derivative (PID) controller75, 76, a model predictive controller (MPC)77, controller based on fuzzy logic78, or a combination of MPC and PID for insulin and glucagon co-delivery79. Dual-hormone or bi-hormonal systems deliver both insulin and glucagon, or another hormone80. Most systems adopt the hybrid approach characterised by manual administration of prandial boluses to mitigate absorption delay of subcutaneously administered rapid-acting insulin.

According to a meta-analysis of 24 RCTs comparing artificial pancreas systems  with control therapy, either conventional pump therapy or sensor-augmented pump therapy, in outpatient settings,81 closed-loop therapy is associated with an increased percentage of time sensor glucose is within the normoglycaemic range, reduced hypoglycaemia while modestly reducing HbA1c.82, 83

With the approval of the first hybrid closed-loop system (MiniMed 670G pump, Medtronic) by the FDA in September 2016 based on a safety study37, 84 and its market introduction in the USA in early 2017, single hormone closed-loop systems have entered the mainstream clinical practice. Further tuning and refinements of the first generation of artificial pancreas systems are expected. Ultra-rapid insulin analogues such as faster insulin aspart, adjunctive therapies using, for example, pramlintide, an SGLT1 or SGLT1&2 inhibitor,85 and inhaled insulin, or ancillary technologies such as site-warming might help to address issues of exercise-induced hypoglycaemia and postprandial hyperglycaemia.

The interest in the artificial pancreas approach is underpinned by do-it-yourself artificial pancreas systems developed and utilised by a small but vocal community of people believing strongly in the potential of the closed-loop approach (e.g. ‘OpenAPS’86, 87 or ‘Loop’)

Data management and technology enabled self-management

Alongside the developments in insulin delivery and glucose monitoring termed as ‘proximal technologies’, advances have been made in the field of ‘distal technologies’ comprising devices and technologies used for communication, education, intervention and remote provision of services. Distal diabetes technologies including telehealth, mobile health applications, game-based support, social platforms, and patient portals were reviewed elsewhere.88 In the context of our review we focus on cloud-upload technologies and applications helping people with type 1 diabetes and healthcare professionals to manage intricate data created by proximal technologies.

Data download

Downloading data from devices such as pumps and CGM monitors to computers and ultimately the cloud enables the user to review summary statistic and to visualise glucose pattern including the ambulatory glucose profile.89 Data are used to assist health care professionals, people with type 1 diabetes, and carergivers to optimise therapy regimens, help users to understand their individual patterns and support self-management. Routinely downloading and reviewing glycaemic data is associated with a significantly lower HbA1c.90 However, only a small percentage of people with type 1 diabetes and/or caregivers routinely download and review.90, 91 Barriers to downloading have not been systematically evaluated but might include inconvenience, difficulties with the software and hardware, and lack of training to interpret the data.

Software tools for downloading devices and data review can be health care professional- or user-operated. Features typically include summary tables, charts such as repeating patterns, blood glucose trends, ambulatory glucose report89 and reviews of pumps and CGM settings. Diabetes device companies offer proprietary tools, e.g. Medtronic Carelink, Dexcom Clarity, Abbott CoPilot, Roche AccuCheck 360, and LifeScan OneTouch, but data aggregator portals exist such as Glooko (now merged with Diasend). Many platforms allow users to link their personal accounts to healthcare provider’s professional accounts for sharing data and remote review enabling more frequent dosing adjustments and faster clinical interventions.

Remote monitoring

CGM devices may send data continuously to cloud, e.g. Dexcom G5 Mobile App, LibreLink App, and Minimed Connect. If enabled, third parties such as partners and caregivers can view users’ CGM traces and receive low glucose or other alerts on their own smartphone. Benefits of remote monitoring in daily living, apart from increased convenience, are yet to be formally evaluated in an RCT. Of note, the development of remote monitoring features was preceded and stimulated by collaborative parent-led Nightscout project.92, 93

Mobile diabetes applications

The use of mobile health applications is increasingly popular, as illustrated by the number of diabetes apps.94 Diabetes apps may improve diabetes self-management offering a wide spectrum of features and activities ranging from simple logs and dosing reminders,95, 96 to bolus calculators and carbohydrate counting, and providing incentives to bolus and peer support.97 Apps like Bant98, Glooko, mySugr, One Drop, and Tidepool99 enable users to keep a detailed diabetes log on their phones. While recent meta-analyses and systematic reviews suggest mobile apps and app-based interventions to support diabetes self-management may improve HbA1c in type 2 diabetes, evidence in type 1 diabetes is limited.100, 101 Hence, providing evidence-based recommendations on app use is difficult.

Clinical practice recommendations

The American Diabetes Association (ADA)102, American Association of Clinical Endocrinologists and American College of Endocrinology (AACE/CE)103, 104, Endocrine Society (ES)105, 106, International Society of Pediatric and Adolescent Diabetes (ISPAD)56, 107, and National Institute for Health and Care Excellence (NICE)108-111 provide recommendations and regular updates regarding clinical indications for use of insulin pump therapy and continuous glucose monitoring. Table 4 summarises these guidelines. Latest developments including flash glucose monitoring, predictive low glucose suspension, and hybrid closed-loop systems, as well as use of mobile apps and non-adjunctive use of CGM have yet to be included, but position statements and practical guidelines on their use are under development.112-116

Ongoing research and further developments

Decision support systems

Traditionally, insulin delivery settings such as insulin-to-carbohydrate ratios, correction factors as well as long acting insulin doses or basal insulin rates were adjusted by health care professionals at follow up clinic visits based on reviews of blood glucose diaries, glucose meter logs and insulin pump downloads. Nowadays, insulin dose self-adjustment is an essential part of intensified insulin therapy and a core element of structured education programmes for people with type 1 diabetes. However, tuning of insulin dosing parameters beyond the simple arithmetic of calculating insulin dosing for meals can be challenging. Multiple groups are developing automated decision support systems by algorithmically optimising dosing recommendations.117-119 Preliminary results look promising and while most systems are still predominantly being used in the research setting, regulatory approval is forthcoming for insulin-dosing decision support systems in type 1 and type 2 diabetes.

Intraperitoneal insulin delivery

From a physiological standpoint, intraperitoneal insulin delivery is attractive due to a faster insulin absorption and first-pass hepatic metabolism. Although implantable pump therapy has been available for over three decades and a potential amelioration of metabolic and endocrine dysregulation in type 1 diabetes has been documented, its use is limited.120 In a closed-loop setting, a proof of concept non-randomized pilot trial compared fully-automated closed-loop delivering intraperitoneal insulin via DiaPort (Roche Diagnostics, Mannheim, Germany) to that with subcutaneous insulin delivery.121 However, the use of intraperitoneal delivery appears limited to people with life-limiting issues related to subcutaneous insulin delivery.

Single port insulin delivery and glucose sensing

Feasibility studies demonstrated that glucose sensing at the site of insulin delivery is feasible.122 ‘Single-port’ devices are being developed, which combine sensor glucose measurements with insulin infusion cannula into a single subcutaneous insulin infusion set intended to simplify device insertion and site management.123, 124 This would be of particular benefits to children with limited amount of skin areas for placement of sensor/insulin cannula but discordance in insertion patterns, cannula is replaced every 2-3 days whereas sensor every 7-14 days, and lower accuracy of sensor during the first day after insertion remains a limiting factor.

Bi-hormonal artificial pancreas

Bi-hormonal or dual hormone closed-loop systems deliver glucagon or other hormones in addition to insulin.80, 125 Glucagon delivery is triggered when hypoglycaemia is impending or predicted. Two approaches have been adopted in terms of the underlying insulin delivery, 80, 125 one tunes insulin delivery in a similar fashion to that of a single hormone closed-loop system, adding glucagon to reduce the residual risk of hypoglycaemia.126 The other approach delivers insulin aggressively and counteracts insulin over-delivery with glucagon. 127

Short-term studies published to date have conﬁrmed reduction in hypoglycaemia with the use of glucagon81. However, such systems have increased complexity and an intricate developmental pathway. Efforts are underway to develop dual-chamber pumps and room-temperature stable glucagon for use in bi-hormonal closed-loop systems128. Long-term data are needed to assess safety and tolerability associated with chronic subcutaneous delivery of glucagon.

Apart from glucagon, other adjunctive therapies including pramlitide and glucagon-like peptide-1 (GLP-1) to suppress post-prandial hyperglucagonaemia and associated hyperglycaemia have been evaluated in combination with close-loop insulin delivery in people with type 1 diabetes in research facility settings129-131.

Bio-artificial pancreas and smart insulins

Apart from algorithm-driven glucose responsive insulin delivery, biological approaches are explored including encapsulated islets,132 glucose responsive polymer encapsulation of insulin, and molecular modification of insulin.133

Islets with their inherent capability of glucose-responsive insulin release are the main focus of the bio-artificial pancreas. To circumvent the need for immunosuppression, islets can be encapsulated within a biocompatible semi-permeable matrix which allows the passage of small molecules such as glucose, insulin and oxygen, but prevent the entry of immune cells and antibodies.132 Intravascular devices contain artificial capillaries and are connected to the host systemic circulation via vascular anastomoses such as deep femoral artery or venous fragment of an arteriovenous anastomosis in the forearm, creating an intravascular shunt. Extravascular devices, sub-classified into micro-capsular and macro-capsular devices do not require creation of intravascular shunts and could be transplanted into different sites such as peritoneum, subcutaneous adipose tissue, or renal capsule.132 Limited data from human trials utilising encapsulated islets xeno- or allografts show a transient reduction in total daily insulin requirements including anecdotal reports of transient insulin independence.134-136 With the rise of pancreatic endocrine cells derived from human embryonic stem cells, pig islets and xenotransplantation have become less favourable.132

In polymer-based smart insulin systems, insulin is encapsulated within glucose-responsive polymeric matrix-based vesicles or hydrogels embedded in large implants, transdermal patches, or microparticles or nanoparticles133. Such matrices are compact during hypoglycaemia or normoglycaemia. Insulin is released in response to rising glucose levels and consecutive structural transformations of the matrices promoting the release of sequestered insulin. Advanced candidates have demonstrated their potential in pre-clinical in vitro studies and animal trials.133 No human data are available.

Other approaches focus on molecular modification of insulin which involves the introduction of a glucose-sensitive motif to the insulin molecule or its formulation, thus promoting an intrinsic glucose responsive activity.133 Present candidate technologies rely on sequestration of active insulin hormone within the subcutaneous space or within the bloodstream (as inactive complexes) with enhanced release or activation only during hyperglycaemia. At present, such strategies remain in the early stages of development.

Conclusions

Diabetes technology is enhancing care and management of people with type 1 diabetes. Innovations in glucose sensing and insulin delivery have reduced burden of self-care and facilitated improved outcomes. A significant milestone has been achieved translating research into clinical use of algorithm-driven glucose responsive insulin delivery in the form of a hybrid artificial pancreas system. In the coming decade, advanced closed-loop systems with add-on data management features will become the standard of care for people with type 1 diabetes across all age groups, while the bio-artificial pancreas and ‘smart’ insulin strategies will take considerably longer to demonstrate safety and efficacy in humans.

Acknowledgments Duality of interest: RH reports having received speaker honoraria from Eli Lilly, Novo Nordisk, and Astra Zeneca, serving on advisory panel for Eli Lilly and Novo Nordisk, receiving license fees from BBraun and Medtronic; having served as a consultant to BBraun, and patents and patent applications related to closed-loop. MT having received speaker honoraria from Medtronic and Novo Nordisk

Funding/support:  Supported by National Institute of Health Research Cambridge Biomedical Research Centre, Efficacy and Mechanism Evaluation National Institute for Health Research (#14/23/09), The Leona M. & Harry B. Helmsley Charitable Trust (#2016PG-T1D045), Wellcome Strategic Award (100574/Z/12/Z), JDRF (#2-SRA-2014-256-M-R), National Institute of Diabetes and Digestive and Kidney Diseases (1UC4DK108520-01), and Diabetes UK (#14/0004878).

Contribution statement: Both authors were responsible for drafting the article and revising it critically for important intellectual content. Both authors approved the version published.

Table 1. Evidence supporting clinical use of pump therapy, continuous glucose monitoring, and glucose responsive insulin delivery in adults with type 1 diabetes

|  | Findings (level of evidence*) | Key references | Highest level of evidence |
| --- | --- | --- | --- |
| Insulin pump therapy | - Reduction in HbA1c compared to multiple daily injection therapy (A) - Comparable to reduced risk of severe hypoglycaemia (A) - Lower insulin requirements (A) - Improved quality of life and treatment satisfaction (A) - Reduced cardiovascular mortality (C) | 14-18 | Systematic reviews/meta-analyses of RCTs |
| Continuous glucose monitoring |  |  |  |
| - Flash glucose monitoring | - Reduction of non-severe hypoglycaemia (A, C) - Improvement in time in range (A) - Improvement in glucose variability (A) - Reduction in HbA1c (A, C) - Improved quality of life/ user satisfaction (C) | 60, 62-64 | One RCT |
| - Real-time continuous glucose monitoring | - Reduction in HbA1c (A) - Reduction of moderate to severe hypoglycaemia (A) - Reduction of time spent in hyperglycaemia (A) - Improvement in quality of life (A) - Benefits of CGM are seen irrespective of insulin delivery method (pump or pen) but are conditioned on high regular sensor usage (A) | 15, 44-46, 48, 50, 51, 54 | Systematic reviews/meta-analyses of RCTs |
| Glucose response insulin delivery |  |  |  |
| - Threshold-based suspension | - Reduced risk of hypoglycaemia in hypo-prone subjects, particularly overnight (A) - No apparent loss in overall glucose control (A) | 66-68 70 | RCTs |
| - Predictive low glucose suspension | - Further reduction of number and duration of diurnal and nocturnal hypoglycaemic events (A) | 72 | RCTs |
| - Hybrid single hormone closed-loop | - Safe use in outpatient settings (A) - Increased time in target glucose range (A) - Reduced time in hypoglycaemia (A) - Reduced time in hyperglycaemia (A) - Modest reduction in HbA1c (A) | 81-84 | Systematic review/meta-analysis |

* based on ADA evidence grading system for clinical recommendations102

Table 2.  Evidence supporting clinical use of pump therapy, continuous glucose monitoring, and glucose responsive insulin delivery in children and adolescents with type 1 diabetes

|  | Findings (level of evidence*) | Key references | Highest level of evidence |
| --- | --- | --- | --- |
| Insulin pump therapy | - Similar to lower HbA1c compared to multiple daily injection therapy; inconclusive benefit in younger children (A) - Inconclusive impact on severe hypoglycaemia - Improved quality of life and treatment satisfaction (A) - Similar risk of diabetic ketoacidosis (A) - Lower insulin requirements (A) | 14-17, 19 | Systematic reviews/meta-analyses of RCTs |
| Continuous glucose monitoring |  |  |  |
| - Flash glucose monitoring | - Similar accuracy as real-time CGM requiring capillary glucose calibrations (C) - No data on clinical effectiveness available | 61 | Non-controlled observational study |
| - Real-time CGM | - Inconclusive impact on HbA1c and hypoglycaemia; benefit conditioned on high regular use (A) | 15, 46, 47, 49, 51, 53 | Systematic reviews/meta-analyses |
| Glucose response insulin delivery |  |  |  |
| - Threshold-based suspension | - Reduced risk of hypoglycaemia, particularly overnight (A) - No apparent loss in overall glucose control (A) | 66, 67, 69 | RCTs |
| - Predictive low glucose suspension | - Further reduction of number and duration of diurnal and nocturnal hypoglycaemic events (A) | 72-74 | RCTs |
| - Hybrid single hormone closed-loop | - Safe use in outpatient settings (A) - Increased time in target glucose range (A) - Reduced time in hypoglycaemia (A) - Reduced time in hyperglycaemia (A) | 37, 81, 82 | Systematic review/meta-analysis |

* based on ADA evidence grading system for clinical recommendations102

Table 3.  ADA evidence-grading system for clinical practice recommendations102

| Level of Evidence | Description |
| --- | --- |
| A | Clear evidence from well-conducted, generalizable randomized controlled trials that are adequately powered, including - Evidence from a well-conducted multicentre trial - Evidence from a meta-analysis that incorporated quality ratings in the analysis Compelling nonexperimental evidence, i.e., “all or none” rule developed by the Centre for Evidence-Based Medicine at the University of Oxford Supportive evidence from well-conducted randomized controlled trials that are adequately powered, including - Evidence from a well-conducted trial at one or more institutions - Evidence from a meta-analysis that incorporated quality ratings in the analysis |
| B | Supportive evidence from well-conducted cohort studies - Evidence from a well-conducted prospective cohort study or registry - Evidence from a well-conducted meta-analysis of cohort studies Supportive evidence from a well-conducted case-control study |
| C | Supportive evidence from poorly controlled or uncontrolled studies - Evidence from randomized clinical trials with one or more major or three or more minor methodological flaws that could invalidate the results - Evidence from observational studies with high potential for bias (such as case series with comparison with historical controls) - Evidence from case series or case reports Conflicting evidence with the weight of evidence supporting the recommendation |
| E | Expert consensus or clinical experience |

Table 4.  Professional society guidelines for use of insulin pump therapy and continuous glucose monitoring in type 1 diabetes (ADA102, AACE/ACE103, 104, ES105, 106, ISPAD56, 107, NICE108-111)

|  | Children and adolescents |  | Adults |  |
| --- | --- | --- | --- | --- |
|  | Insulin pump therapy | Continuous glucose monitoring | Insulin pump therapy | Continuous glucose monitoring |
| Indication/Criterion |  |  |  |  |
| HbA1c | - Recommended if HbA1c is persistently above the individual goal (ISPAD) - Elevated HbA1c levels on injection therapy (AACE/ACE) | - In children and adolescents who have HbA1c levels <7.0% [58mmol/mol] as it will assist in maintaining target HbA1c levels while limiting the risk of hypoglycaemia. - In children and adolescents with HbA1c ≥ 7.0% [58mmol/mol] and who are able to use the devices on a nearly daily basis (ES) | - Patients with type 1 diabetes who do not reach glycaemic goals despite adherence to maximum MDI (AACE/ACE) - HbA1c levels have remained high (> 8.5% [69 mmol/mol] or above) on MDI (NICE). | - Patients with HbA1c levels above and below target (ES) - Hyperglycaemia (HbA1c level of 75 mmol/mol [9%] or higher) that persists despite testing at least 10 times a day (NICE) - Continue real-time CGM only if HbA1c can be sustained at or below 7% (53 mmol/mol) and/or there has been a fall in HbA1c of 2.5% (27 mmol/mol) or more (NICE). |
| Hypoglycaemia | - If hypoglycaemia is a major problem (ISPAD) - Frequent and severe hypoglycaemia (AACE/ACE) | - CGM recommended for patients with impaired awareness of hypoglycaemia (AACE, ADA, ES, ISPAD, NICE) - History of severe hypoglycaemia (AACE), - Frequent severe hypoglycaemia (NICE) - Frequent hypoglycaemia (ADA), nocturnal hypoglycaemia (ES) - Inability to recognise, or communicate about, symptoms of hypoglycaemia (e.g. because of cognitive or neurological disabilities) (NICE) | - Attempts to achieve target haemoglobin A1c (HbA1c) levels with MDI result in the person experiencing disabling hypoglycaemia (NICE) - Frequent severe hypoglycaemia and/or hypoglycaemia unawareness (AACE/ACE) | - Hypoglycaemia unawareness (ADA, NICE) - Frequent hypoglycaemic episodes (ADA) - Frequent (more than 2 episodes a week) asymptomatic hypoglycaemia that is causing problems with daily activities (NICE) - More than 1 episode a year of severe hypoglycaemia with no obviously preventable precipitating cause (NICE) - Extreme fear of hypoglycaemia (NICE) |
| Hyperglycaemia, ketonaemia, glucose variability | - Patients with pronounced dawn phenomenon (AACE/ACE) - Ketosis prone patients (AACE/ACE) - Widely fluctuating glucose levels (AACE/ACE) | - Intermittent use in children and young people who continue to have hyperglycaemia despite insulin adjustment and additional support | - Patients who have very labile diabetes (erratic glucose control including recurrent DKA) (AACE/ACE) - Significant dawn phenomenon (AACE/ACE) |  |
| Quality of life/ psychosocial aspects | - If quality of life needs be improved (ISPAD) - If current treatment regimen compromises lifestyle (AACE/ACE) - Children with pronounced needle phobia (AACE/ACE) - Treatment option for children younger than 12 years provided that MDI therapy is considered to be impractical or inappropriate (NICE) |  | - Fear of hypoglycaemia (NICE) |  |
| Education/training, adherence, follow up | - Ideal paediatric candidates are those with motivated families who are committed to monitoring blood glucose ≥4 times/day and have a working understanding of basic diabetes management (AACE/ACE) | - In patients who are able to use the devices on a nearly daily basis (ES) - Robust diabetes education, training and support are required for optimal CGM implementation and ongoing use (ADA) - CGM users must know the basics of sensor insertion, calibration and real-time data interpretation (AACE/ACE) | - Currently performing ≥4 insulin injections and ≥4 self-monitored blood glucose measurements daily (AACE/ACE) - Motivated to achieve optimal blood glucose control (AACE/ACE) - Regular follow up with specialist team (AACE/ACE) | - When prescribing CGM, robust diabetes education, training, and support are required for optimal CGM implementation and ongoing use - In patients who are willing and able to use these devices on a nearly daily basis (ES). - CGM to be considered in adults who are willing to commit to using it at least 70% of the time and to calibrate it as needed (NICE) |
| Special populations | - For patients with microvascular complications and/or risk factors for macrovascular complications (AACE/ACE). - Very young children (AACE/ACE) - Athletes (AACE/ACE) - Adolescents with eating disorders (AACE/ACE) - Pregnant teens (AACE/ACE) | - CGM to be considered in neonates, infants and pre-school children (NICE) - CGM to be considered in children and young people who have comorbidities or who are receiving treatments that can make blood glucose control difficult (NICE) - CGM to be considered in children and young people who undertake high levels of physical activity (NICE) - No recommendations about the use of CGM in children less than 8 years of age (ES) | - Preconception (AACE/ACE) - Pregnancy (AACE/ACE) - competitive athletes (AACE/ACE) - Patients with extreme insulin sensitivity (AACE/ACE) |  |

Figure 1. Currently available diabetes technologies in type 1 diabetes

References

1.	Bergenstal RM. Continuous glucose monitoring: transforming diabetes management step by step. Lancet 2018 Feb.

2.	International Diabetes Federation. Diabetes Atlas (7th edition).  2015.

3.	Diabetes Control and Complications Trial Research Group. The effect of intensive treatment of diabetes on the development and progression of long-term complications in insulin-dependent diabetes mellitus. The Diabetes Control and Complications Trial Research Group. N Engl J Med 1993 Sep 30;329(14):977-86.

4.	Paton JS, Wilson M, Ireland JT, Reith SB. Convenient pocket insulin syringe. Lancet 1981 Jan;1(8213):189-90.

5.	Pickup JC, Keen H, Parsons JA, Alberti KG. Continuous subcutaneous insulin infusion: an approach to achieving normoglycaemia. Br Med J 1978 Jan;1(6107):204-7.

6.	Renard E. Insulin pump use in Europe. Diabetes Technol Ther 2010 Jun;12 Suppl 1:S29-32.

7.	Szypowska A, Schwandt A, Svensson J, Shalitin S, Cardona-Hernandez R, Forsander G, et al. Insulin pump therapy in children with type 1 diabetes: analysis of data from the SWEET registry. Pediatr Diabetes 2016 10;17 Suppl 23:38-45.

8.	Sherr JL, Hermann JM, Campbell F, Foster NC, Hofer SE, Allgrove J, et al. Use of insulin pump therapy in children and adolescents with type 1 diabetes and its impact on metabolic control: comparison of results from three large, transatlantic paediatric registries. Diabetologia 2016 Jan;59(1):87-91.

9.	Zisser HC. The OmniPod Insulin Management System: the latest innovation in insulin pump therapy. Diabetes Ther 2010 Aug;1(1):10-24.

10.	Leelarathna L, Roberts SA, Hindle A, Markakis K, Alam T, Chapman A, et al. Comparison of different insulin pump makes under routine care conditions in adults with Type 1 diabetes. Diabet Med 2017 Oct;34(10):1372-79.

11.	Ramotowska A, Golicki D, Dzygalo K, Szypowska A. The effect of using the insulin pump bolus calculator compared to standard insulin dosage calculations in patients with type 1 diabetes mellitus - systematic review. Exp Clin Endocrinol Diabetes 2013 May;121(5):248-54.

12.	Jones SM, Quarry JL, Caldwell-McMillan M, Mauger DT, Gabbay RA. Optimal insulin pump dosing and postprandial glycemia following a pizza meal using the continuous glucose monitoring system. Diabetes Technol Ther 2005 Apr;7(2):233-40.

13.	O'Connell MA, Gilbertson HR, Donath SM, Cameron FJ. Optimizing postprandial glycemia in pediatric patients with type 1 diabetes using insulin pump therapy: impact of glycemic index and prandial bolus type. Diabetes Care 2008 Aug;31(8):1491-5.

14.	Pickup JC, Sutton AJ. Severe hypoglycaemia and glycaemic control in Type 1 diabetes: meta-analysis of multiple daily insulin injections compared with continuous subcutaneous insulin infusion. Diabet Med 2008 Jul;25(7):765-74.

15.	Yeh HC, Brown TT, Maruthur N, Ranasinghe P, Berger Z, Suh YD, et al. Comparative effectiveness and safety of methods of insulin delivery and glucose monitoring for diabetes mellitus: a systematic review and meta-analysis. Ann Intern Med 2012 Sep 4;157(5):336-47.

16.	Jeitler K, Horvath K, Berghold A, Gratzer TW, Neeser K, Pieber TR, et al. Continuous subcutaneous insulin infusion versus multiple daily insulin injections in patients with diabetes mellitus: systematic review and meta-analysis. Diabetologia 2008 Jun;51(6):941-51.

17.	Fatourechi MM, Kudva YC, Murad MH, Elamin MB, Tabini CC, Montori VM. Clinical review: Hypoglycemia with intensive insulin therapy: a systematic review and meta-analyses of randomized trials of continuous subcutaneous insulin infusion versus multiple daily injections. J Clin Endocrinol Metab 2009 Mar;94(3):729-40.

18.	Monami M, Lamanna C, Marchionni N, Mannucci E. Continuous subcutaneous insulin infusion versus multiple daily insulin injections in type 1 diabetes: a meta-analysis. Acta Diabetol 2010 Dec;47 Suppl 1:77-81.

19.	Pańkowska E, Błazik M, Dziechciarz P, Szypowska A, Szajewska H. Continuous subcutaneous insulin infusion vs. multiple daily injections in children with type 1 diabetes: a systematic review and meta-analysis of randomized control trials. Pediatr Diabetes 2009 Feb;10(1):52-8.

20.	Beato-Víbora P, Yeoh E, Rogers H, Hopkins D, Amiel SA, Choudhary P. Sustained benefit of continuous subcutaneous insulin infusion on glycaemic control and hypoglycaemia in adults with Type 1 diabetes. Diabet Med 2015 Nov;32(11):1453-9.

21.	Joubert M, Morera J, Vicente A, Rod A, Parienti JJ, Reznik Y. Cross-sectional survey and retrospective analysis of a large cohort of adults with type 1 diabetes with long-term continuous subcutaneous insulin infusion treatment. J Diabetes Sci Technol 2014 Sep;8(5):1005-10.

22.	Orr CJ, Hopman W, Yen JL, Houlden RL. Long-term efficacy of insulin pump therapy on glycemic control in adults with type 1 diabetes mellitus. Diabetes Technol Ther 2015 Jan;17(1):49-54.

23.	Johnson SR, Cooper MN, Jones TW, Davis EA. Long-term outcome of insulin pump therapy in children with type 1 diabetes assessed in a large population-based case-control study. Diabetologia 2013 Aug;56(11):2392-400.

24.	Steineck I, Cederholm J, Eliasson B, Rawshani A, Eeg-Olofsson K, Svensson AM, et al. Insulin pump therapy, multiple daily injections, and cardiovascular mortality in 18,168 people with type 1 diabetes: observational study. BMJ 2015;350:h3234.

25.	Karges B, Schwandt A, Heidtmann B, Kordonouri O, Binder E, Schierloh U, et al. Association of insulin pump therapy vs insulin injection therapy with severe hypoglycemia, ketoacidosis, and glycemic control among children, adolescents, and young adults with type 1 diabetes. JAMA 2017 10;318(14):1358-66.

26.	Lin MH, Connor CG, Ruedy KJ, Beck RW, Kollman C, Buckingham B, et al. Race, socioeconomic status, and treatment center are associated with insulin pump therapy in youth in the first year following diagnosis of type 1 diabetes. Diabetes Technol Ther 2013 Nov;15(11):929-34.

27.	Ziegler R, Heidtmann B, Hilgard D, Hofer S, Rosenbauer J, Holl R, et al. Frequency of SMBG correlates with HbA1c and acute complications in children and adolescents with type 1 diabetes. Pediatr Diabetes 2011 Feb;12(1):11-7.

28.	Miller KM, Beck RW, Bergenstal RM, Goland RS, Haller MJ, McGill JB, et al. Evidence of a strong association between frequency of self-monitoring of blood glucose and hemoglobin A1c levels in T1D exchange clinic registry participants. Diabetes Care 2013 Jul;36(7):2009-14.

29.	Ziegler R, Cavan DA, Cranston I, Barnard K, Ryder J, Vogel C, et al. Use of an insulin bolus advisor improves glycemic control in multiple daily insulin injection (MDI) therapy patients with suboptimal glycemic control: First results from the ABACUS trial. Diabetes Care 2013 Nov;36(11):3613-9.

30.	Vallejo Mora MDR, Carreira M, Anarte MT, Linares F, Olveira G, González Romero S. Bolus calculator reduces hypoglycemia in the short term and fear of hypoglycemia in the long term in subjects with type 1 diabetes (CBMDI study). Diabetes Technol Ther 2017 Jul;19(7):402-09.

31.	Vallejo-Mora MD, Carreira-Soler M, Linares-Parrado F, Olveira G, Rojo-Martínez G, Domínguez-López M, et al. The Calculating Boluses on Multiple Daily Injections (CBMDI) study: A randomized controlled trial on the effect on metabolic control of adding a bolus calculator to multiple daily injections in people with type 1 diabetes. J Diabetes 2017 Jan;9(1):24-33.

32.	Bailey T, Bode BW, Christiansen MP, Klaff LJ, Alva S. The performance and usability of a factory-calibrated flash glucose monitoring system. Diabetes Technol Ther 2015 Jul;17(11):787-94.

33.	Kropff J, Choudhary P, Neupane S, Barnard K, Bain SC, Kapitza C, et al. Accuracy and Longevity of an Implantable Continuous Glucose Sensor in the PRECISE Study: A 180-Day, Prospective, Multicenter, Pivotal Trial. Diabetes Care 2017 Jan;40(1):63-68.

34.	DeSalvo D, Miller K, Hermann J, Maahs D, Hofer S, Clements M, et al. Continuous Glucose Monitoring (CGM) and Glycemic Control Among Youth with Type 1 Diabetes (T1D): International comparison from the T1D Exchange (T1DX) and the DPV Initiative.  43rd Annual Conference of the International Society for Pediatric and Adolescent Diabetes. Innsbruck, Austria 2017.

35.	Bailey TS, Chang A, Christiansen M. Clinical accuracy of a continuous glucose monitoring system with an advanced algorithm. J Diabetes Sci Technol 2015 Mar;9(2):209-14.

36.	Laffel L. Improved accuracy of continuous glucose monitoring systems in pediatric patients with diabetes mellitus: Results from two studies. Diabetes Technol Ther 2016 Feb;18 Suppl 2:S223-33.

37.	Garg SK, Weinzimer SA, Tamborlane WV, Buckingham BA, Bode BW, Bailey TS, et al. Glucose outcomes with the in-home use of a hybrid closed-loop insulin delivery system in adolescents and adults with type 1 diabetes. Diabetes Technol Ther 2017 Jan.

38.	Kropff J, Bruttomesso D, Doll W, Farret A, Galasso S, Luijf YM, et al. Accuracy of two continuous glucose monitoring systems: a head-to-head comparison under clinical research centre and daily life conditions. Diabetes Obes Metab 2015 Apr;17(4):343-9.

39.	Pleus S, Schoemaker M, Morgenstern K, Schmelzeisen-Redeker G, Haug C, Link M, et al. Rate-of-change dependence of the performance of two CGM systems during induced glucose swings. J Diabetes Sci Technol 2015 Jul;9(4):801-7.

40.	Kovatchev BP, Patek SD, Ortiz EA, Breton MD. Assessing sensor accuracy for non-adjunct use of continuous glucose monitoring. Diabetes Technol Ther 2015 Mar;17(3):177-86.

41.	Aleppo G, Ruedy KJ, Riddlesworth TD, Kruger DF, Peters AL, Hirsch I, et al. REPLACE-BG: A randomized trial comparing continuous glucose monitoring with and without routine blood glucose monitoring in adults with well-controlled type 1 diabetes. Diabetes Care 2017 Apr;40(4):538-45.

42.	Agiostratidou G, Anhalt H, Ball D, Blonde L, Gourgari E, Harriman KN, et al. Standardizing Clinically Meaningful Outcome Measures Beyond HbA1c for Type 1 Diabetes: A Consensus Report of the American Association of Clinical Endocrinologists, the American Association of Diabetes Educators, the American Diabetes Association, the Endocrine Society, JDRF International, The Leona M. and Harry B. Helmsley Charitable Trust, the Pediatric Endocrine Society, and the T1D Exchange. Diabetes Care 2017 Dec;40(12):1622-30.

43.	Danne T, Nimri R, Battelino T, Bergenstal RM, Close KL, DeVries JH, et al. International Consensus on Use of Continuous Glucose Monitoring. Diabetes Care 2017 Dec;40(12):1631-40.

44.	Juvenile Diabetes Research Foundation Continuous Glucose Monitoring Study G, Tamborlane WV, Beck RW, Bode BW, Buckingham B, Chase HP, et al. Continuous glucose monitoring and intensive treatment of type 1 diabetes. N Engl J Med 2008 Oct 2;359(14):1464-76.

45.	Pickup JC, Freeman SC, Sutton AJ. Glycaemic control in type 1 diabetes during real time continuous glucose monitoring compared with self monitoring of blood glucose: meta-analysis of randomised controlled trials using individual patient data. BMJ 2011;343:d3805.

46.	Langendam M, Luijf YM, Hooft L, Devries JH, Mudde AH, Scholten RJ. Continuous glucose monitoring systems for type 1 diabetes mellitus. The Cochrane database of systematic reviews 2012 Jan 18;1:CD008101.

47.	Szypowska A, Ramotowska A, Dzygalo K, Golicki D. Beneficial effect of real-time continuous glucose monitoring system on glycemic control in type 1 diabetic patients: systematic review and meta-analysis of randomized trials. Eur J Endocrinol 2012 Apr;166(4):567-74.

48.	Hoeks LB, Greven WL, de Valk HW. Real-time continuous glucose monitoring system for treatment of diabetes: a systematic review. Diabet Med 2011 Apr;28(4):386-94.

49.	Golicki DT, Golicka D, Groele L, Pankowska E. Continuous glucose monitoring system in children with type 1 diabetes mellitus: a systematic review and meta-analysis. Diabetologia 2008 Feb;51(2):233-40.

50.	Juvenile Diabetes Research Foundation Continuous Glucose Monitoring Study Group. Effectiveness of continuous glucose monitoring in a clinical care environment: evidence from the Juvenile Diabetes Research Foundation continuous glucose monitoring (JDRF-CGM) trial. Diabetes Care 2010 Jan;33(1):17-22.

51.	Bergenstal RM, Tamborlane WV, Ahmann A, Buse JB, Dailey G, Davis SN, et al. Effectiveness of sensor-augmented insulin-pump therapy in type 1 diabetes. N Engl J Med 2010 Jul 22;363(4):311-20.

52.	Bergenstal RM, Tamborlane WV, Ahmann A, Buse JB, Dailey G, Davis SN, et al. Sensor-augmented pump therapy for A1C reduction (STAR 3) study: results from the 6-month continuation phase. Diabetes Care 2011 Nov;34(11):2403-5.

53.	Battelino T, Phillip M, Bratina N, Nimri R, Oskarsson P, Bolinder J. Effect of continuous glucose monitoring on hypoglycemia in type 1 diabetes. Diabetes Care 2011 Apr;34(4):795-800.

54.	Battelino T, Conget I, Olsen B, Schutz-Fuhrmann I, Hommel E, Hoogma R, et al. The use and efficacy of continuous glucose monitoring in type 1 diabetes treated with insulin pump therapy: a randomised controlled trial. Diabetologia 2012 Dec;55(12):3155-62.

55.	Beck RW, Riddlesworth T, Ruedy K, Ahmann A, Bergenstal R, Haller S, et al. Effect of continuous glucose monitoring on glycemic control in adults with type 1 diabetes using insulin injections: The DIAMOND randomized clinical trial. JAMA 2017 01;317(4):371-78.

56.	Rewers MJ, Pillay K, de Beaufort C, Craig ME, Hanas R, Acerini CL, et al. ISPAD clinical practice consensus guidelines 2014. Assessment and monitoring of glycemic control in children and adolescents with diabetes. Pediatr Diabetes 2014 Sep;15 Suppl 20:102-14.

57.	Fonseca VA, Grunberger G, Anhalt H, Bailey TS, Blevins T, Garg SK, et al. Continuous glucose monitoring: A consensus conference of the American Association of Clinical Endocrinologists and American College of Endocrinology. Endocr Pract 2016 Aug;22(8):1008-21.

58.	Foster NC, Miller KM, Tamborlane WV, Bergenstal RM, Beck RW, Network TDEC. Continuous Glucose Monitoring in Patients With Type 1 Diabetes Using Insulin Injections. Diabetes Care 2016 Jun;39(6):e81-2.

59.	Lind M, Polonsky W, Hirsch IB, Heise T, Bolinder J, Dahlqvist S, et al. Continuous glucose monitoring vs conventional therapy for glycemic control in adults with type 1 diabetes treated with multiple daily insulin injections: The GOLD randomized clinical trial. JAMA 2017 01;317(4):379-87.

60.	Bolinder J, Antuna R, Geelhoed-Duijvestijn P, Kröger J, Weitgasser R. Novel glucose-sensing technology and hypoglycaemia in type 1 diabetes: a multicentre, non-masked, randomised controlled trial. Lancet 2016 11;388(10057):2254-63.

61.	Edge J, Acerini C, Campbell F, Hamilton-Shield J, Moudiotis C, Rahman S, et al. An alternative sensor-based method for glucose monitoring in children and young people with diabetes. Arch Dis Child 2017 Jun;102(6):543-49.

62.	Ish-Shalom M, Wainstein J, Raz I, Mosenzon O. Improvement in glucose control in difficult-to-control patients with diabetes using a novel flash glucose monitoring device. J Diabetes Sci Technol 2016 Nov;10(6):1412-13.

63.	Dover AR, Stimson RH, Zammitt NN, Gibb FW. Flash glucose monitoring improves outcomes in a type 1 diabetes clinic. J Diabetes Sci Technol 2017 Mar;11(2):442-43.

64.	McKnight JA, Gibb FW. Flash glucose monitoring is associated with improved glycaemic control but use is largely limited to more affluent people in a UK diabetes centre. Diabet Med 2017 May;34(5):732.

65.	Dunn TC, Xu Y, Hayter G, Ajjan RA. Real-world flash glucose monitoring patterns and associations between self-monitoring frequency and glycaemic measures: A European analysis of over 60 million glucose tests. Diabetes Res Clin Pract 2017 Dec.

66.	Ly TT, Nicholas JA, Retterath A, Lim EM, Davis EA, Jones TW. Effect of sensor-augmented insulin pump therapy and automated insulin suspension vs standard insulin pump therapy on hypoglycemia in patients with type 1 diabetes: a randomized clinical trial. JAMA 2013 Sep;310(12):1240-7.

67.	Bergenstal RM, Klonoff DC, Garg SK, Bode BW, Meredith M, Slover RH, et al. Threshold-based insulin-pump interruption for reduction of hypoglycemia. N Engl J Med 2013 Jul;369(3):224-32.

68.	Weiss R, Garg SK, Bode BW, Bailey TS, Ahmann AJ, Schultz KA, et al. Hypoglycemia reduction and changes in hemoglobin A1c in the ASPIRE in-home study. Diabetes Technol Ther 2015 Aug;17(8):542-7.

69.	Danne T, Kordonouri O, Holder M, Haberland H, Golembowski S, Remus K, et al. Prevention of hypoglycemia by using low glucose suspend function in sensor-augmented pump therapy. Diabetes Technol Ther 2011 Nov;13(11):1129-34.

70.	Choudhary P, Shin J, Wang Y, Evans ML, Hammond PJ, Kerr D, et al. Insulin pump therapy with automated insulin suspension in response to hypoglycemia: reduction in nocturnal hypoglycemia in those at greatest risk. Diabetes Care 2011 Sep;34(9):2023-5.

71.	Agrawal P, Zhong A, Welsh JB, Shah R, Kaufman FR. Retrospective analysis of the real-world use of the threshold suspend feature of sensor-augmented insulin pumps. Diabetes Technol Ther 2015 May;17(5):316-9.

72.	Maahs DM, Calhoun P, Buckingham BA, Chase HP, Hramiak I, Lum J, et al. A randomized trial of a home system to reduce nocturnal hypoglycemia in type 1 diabetes. Diabetes Care 2014 Jul;37(7):1885-91.

73.	Buckingham BA, Raghinaru D, Cameron F, Bequette BW, Chase HP, Maahs DM, et al. Predictive low-glucose insulin suspension reduces duration of nocturnal hypoglycemia in children without increasing ketosis. Diabetes Care 2015 Jul;38(7):1197-204.

74.	Battelino T, Nimri R, Dovc K, Phillip M, Bratina N. Prevention of hypoglycemia with predictive low glucose insulin suspension in children with type 1 diabetes: A randomized controlled trial. Diabetes Care 2017 Mar;40(6):764-70.

75.	Steil GM, Rebrin K, Darwin C, Hariri F, Saad MF. Feasibility of automating insulin delivery for the treatment of type 1 diabetes. Diabetes 2006 Dec;55(12):3344-50.

76.	Weinzimer SA, Steil GM, Swan KL, Dziura J, Kurtz N, Tamborlane WV. Fully automated closed-loop insulin delivery versus semiautomated hybrid control in pediatric patients with type 1 diabetes using an artificial pancreas. Diabetes Care 2008 May;31(5):934-9.

77.	Hovorka R, Allen JM, Elleri D, Chassin LJ, Harris J, Xing D, et al. Manual closed-loop insulin delivery in children and adolescents with type 1 diabetes: a phase 2 randomised crossover trial. Lancet 2010 Feb;375(9716):743-51.

78.	Atlas E, Nimri R, Miller S, Grunberg EA, Phillip M. MD-logic artificial pancreas system: a pilot study in adults with type 1 diabetes. Diabetes Care 2010 May;33(5):1072-6.

79.	El-Khatib FH, Russell SJ, Nathan DM, Sutherlin RG, Damiano ER. A bihormonal closed-loop artificial pancreas for type 1 diabetes. Sci Transl Med 2010 Apr 14;2(27):27ra27.

80.	Bakhtiani PA, Zhao LM, El Youssef J, Castle JR, Ward WK. A review of artificial pancreas technologies with an emphasis on bi-hormonal therapy. Diabetes Obes Metab 2013 Apr;15(12):1065-70.

81.	Weisman A, Bai JW, Cardinez M, Kramer CK, Perkins BA. Effect of artificial pancreas systems on glycaemic control in patients with type 1 diabetes: a systematic review and meta-analysis of outpatient randomised controlled trials. Lancet Diabetes Endocrinol 2017 Jul;5(7):501-12.

82.	Thabit H, Tauschmann M, Allen JM, Leelarathna L, Hartnell S, Wilinska ME, et al. Home use of an artificial beta cell in type 1 diabetes. N Engl J Med 2015 Nov;373(22):2129-40.

83.	Kropff J, Del Favero S, Place J, Toffanin C, Visentin R, Monaro M, et al. 2 month evening and night closed-loop glucose control in patients with type 1 diabetes under free-living conditions: a randomised crossover trial. Lancet Diabetes Endocrinol 2015 Dec;3(12):939-47.

84.	Bergenstal RM, Garg S, Weinzimer SA, Buckingham BA, Bode BW, Tamborlane WV, et al. Safety of a hybrid closed-loop insulin delivery system in patients with type 1 diabetes. JAMA 2016 Oct;316(13):1407-08.

85.	Garg SK, Henry RR, Banks P, Buse JB, Davies MJ, Fulcher GR, et al. Effects of sotagliflozin added to insulin in patients with type 1 diabetes. N Engl J Med 2017 12;377(24):2337-48.

86.	Lewis D. OpenAPS.  2017  [cited 2017 15 November]; Available from: https://openaps.org/

87.	Lewis D, Leibrand S, Community O. Real-world use of open source artificial pancreas systems. J Diabetes Sci Technol 2016 Nov;10(6):1411.

88.	Duke DC, Barry S, Wagner DV, Speight J, Choudhary P, Harris MA. Distal technologies and type 1 diabetes management. Lancet Diabetes Endocrinol 2017 Aug.

89.	Mazze RS, Strock E, Wesley D, Borgman S, Morgan B, Bergenstal R, et al. Characterizing glucose exposure for individuals with normal glucose tolerance using continuous glucose monitoring and ambulatory glucose profile analysis. Diabetes Technol Ther 2008 Jun;10(3):149-59.

90.	Wong JC, Neinstein AB, Spindler M, Adi S. A minority of patients with type 1 diabetes routinely downloads and retrospectively reviews device data. Diabetes Technol Ther 2015 Aug;17(8):555-62.

91.	Beck RW. Downloading diabetes device data: Empowering patients to download at home to achieve better outcomes. Diabetes Technol Ther 2015 Aug;17(8):536-7.

92.	Lee JM, Newman MW, Gebremariam A, Choi P, Lewis D, Nordgren W, et al. Real-world use and self-reported health outcomes of a patient-designed do-it-yourself mobile technology system for diabetes: Lessons for mobile health. Diabetes Technol Ther 2017 Apr;19(4):209-19.

93.	Lee JM, Hirschfeld E, Wedding J. A patient-designed do-it-yourself mobile technology system for diabetes: promise and challenges for a new era in medicine. JAMA 2016 Apr;315(14):1447-8.

94.	Trawley S, Browne JL, Hagger VL, Hendrieckx C, Holmes-Truscott E, Pouwer F, et al. The use of mobile applications among adolescents with type 1 diabetes: Results from Diabetes MILES Youth-Australia. Diabetes Technol Ther 2016 Dec;18(12):813-19.

95.	Frøisland DH, Arsand E, Skårderud F. Improving diabetes care for young people with type 1 diabetes through visual learning on mobile phones: mixed-methods study. J Med Internet Res 2012 Aug;14(4):e111.

96.	Hanauer DA, Wentzell K, Laffel N, Laffel LM. Computerized Automated Reminder Diabetes System (CARDS): e-mail and SMS cell phone text messaging reminders to support diabetes management. Diabetes Technol Ther 2009 Feb;11(2):99-106.

97.	Cafazzo JA, Casselman M, Hamming N, Katzman DK, Palmert MR. Design of an mHealth app for the self-management of adolescent type 1 diabetes: a pilot study. J Med Internet Res 2012 May;14(3):e70.

98.	Goyal S, Nunn CA, Rotondi M, Couperthwaite AB, Reiser S, Simone A, et al. A mobile app for the self-management of type 1 diabetes among adolescents: A randomized controlled trial. JMIR Mhealth Uhealth 2017 Jun;5(6):e82.

99.	Neinstein A, Wong J, Look H, Arbiter B, Quirk K, McCanne S, et al. A case study in open source innovation: developing the Tidepool Platform for interoperability in type 1 diabetes management. J Am Med Inform Assoc 2016 Mar;23(2):324-32.

100.	Hou C, Carter B, Hewitt J, Francisa T, Mayor S. Do mobile phone applications improve glycemic control (HbA1c) in the self-management of diabetes? A systematic review, meta-analysis, and GRADE of 14 randomized trials. Diabetes Care 2016 Nov;39(11):2089-95.

101.	Wu Y, Yao X, Vespasiani G, Nicolucci A, Dong Y, Kwong J, et al. Mobile app-based interventions to support diabetes self-management: A systematic review of randomized controlled trials to identify functions associated with glycemic efficacy. JMIR Mhealth Uhealth 2017 Mar;5(3):e35.

102.	American Diabetes Association. Standards of medical care in diabetes - 2017. Diabetes Care 2017;40(Suppl. 1):S1-135.

103.	Bailey TS, Grunberger G, Bode BW, Handelsman Y, Hirsch IB, Jovanovič L, et al. American Assiciation of Clinical Endocrinologists and American College of Endocrinology 2016 outpatient glucose monitoring consensus statement. Endocr Pract 2016 Feb;22(2):231-61.

104.	Grunberger G, Abelseth JM, Bailey TS, Bode BW, Handelsman Y, Hellman R, et al. Consensus statement by the American Association of Clinical Endocrinologists/American College of Endocrinology insulin pump management task force. Endocr Pract 2014 May;20(5):463-89.

105.	Klonoff DC, Buckingham B, Christiansen JS, Montori VM, Tamborlane WV, Vigersky RA, et al. Continuous glucose monitoring: an Endocrine Society Clinical Practice Guideline. J Clin Endocrinol Metab 2011 Oct;96(10):2968-79.

106.	Peters AL, Ahmann AJ, Battelino T, Evert A, Hirsch IB, Murad MH, et al. Diabetes technology-continuous subcutaneous insulin infusion therapy and continuous glucose monitoring in adults: An Endocrine Society clinical practice guideline. J Clin Endocrinol Metab 2016 Nov;101(11):3922-37.

107.	Danne T, Bangstad HJ, Deeb L, Jarosz-Chobot P, Mungaie L, Saboo B, et al. ISPAD clinical practice consensus guidelines 2014. Insulin treatment in children and adolescents with diabetes. Pediatr Diabetes 2014 Sep;15 Suppl 20:115-34.

108.	National Institute for Health and Care Excellence (NICE). NICE guideline (NG) 18, Diabetes (type 1 and type 2) in children and young people: diagnosis and management.  26 August 2015  [cited; Available from: http://nice.org.uk/guidance/ng18

109.	National Institute for Health and Clinical Excellence (NICE). NICE guideline (NG) 17, Type 1 diabetes in adults: diagnosis and management.  26 August 2015  [cited; Available from: http://nice.org.uk/guidance/ng17

110.	National Institute for Health and Care Excellence (NICE). Technology appraisal guidance (TA151). Continuous subcutaneous insulin infusion for the treatment of diabetes mellitus.  23 July 2008  [cited; Available from: nice.org.uk/guidance/ta151

111.	National Institute for Health and Care Excellence (NICE). Diagnostics guidance (DG21). Integrated sensor-augmented pump therapy systems for managing blood glucose levels in type 1 diabetes (the MiniMed Paradigm Veo system and the Vibe and G4 PLATINUM CGM system).  12 February 2016  [cited; Available from: nice.org.uk/guidance/dg21

112.	Edelman SV. Regulation catches up to reality. J Diabetes Sci Technol 2017 Jan;11(1):160-64.

113.	Pettus J, Edelman SV. Recommendations for using real-time continuous glucose monitoring (rtCGM) data for insulin adjustments in type 1 diabetes. J Diabetes Sci Technol 2017 Jan;11(1):138-47.

114.	Castle JR, Jacobs PG. Nonadjunctive use of continuous glucose monitoring for diabetes treatment decisions. J Diabetes Sci Technol 2016 Sep;10(5):1169-73.

115.	Forlenza GP, Argento NB, Laffel LM. Practical considerations on the use of continuous glucose monitoring in pediatrics and older adults and nonadjunctive use. Diabetes Technol Ther 2017 Jun;19(S3):S13-S20.

116.	National Institute for Health and Care Excellence (NICE). FreeStyle Libre for glucose monitoring. Medtech innovation briefing.  July 2017  [cited; Available from: http://nice.org.uk/guidance/mib110

117.	Wang Y, Zhang J, Zeng F, Wang N, Chen X, Zhang B, et al. "Learning" can improve the blood glucose control performance for type 1 diabetes mellitus. Diabetes Technol Ther 2017 01;19(1):41-48.

118.	Dassau E, Brown SA, Basu A, Pinsker JE, Kudva YC, Gondhalekar R, et al. Adjustment of open-loop settings to improve closed-loop results in type 1 diabetes: A multicenter randomized trial. J Clin Endocrinol Metab 2015 Oct;100(10):3878-86.

119.	Reddy M, Pesl P, Xenou M, Toumazou C, Johnston D, Georgiou P, et al. Clinical safety and feasibility of the advanced bolus calculator for type 1 diabetes based on case-based reasoning: A 6-week nonrandomized single-arm pilot study. Diabetes Technol Ther 2016 08;18(8):487-93.

120.	Bally L, Thabit H, Hovorka R. Finding the right route for insulin delivery - an overview of implantable pump therapy. Expert Opin Drug Deliv 2017 Sep;14(9):1103-11.

121.	Dassau E, Renard E, Place J, Farret A, Pelletier MJ, Lee J, et al. Intraperitoneal insulin delivery provides superior glycaemic regulation to subcutaneous insulin delivery in model predictive control-based fully-automated artificial pancreas in patients with type 1 diabetes: a pilot study. Diabetes Obes Metab 2017 Dec;19(12):1698-705.

122.	Lindpointner S, Korsatko S, Kohler G, Kohler H, Schaller R, Kaidar R, et al. Use of the site of subcutaneous insulin administration for the measurement of glucose in patients with type 1 diabetes. Diabetes Care 2010 Mar;33(3):595-601.

123.	Regittnig W, Lindpointner S, Korsatko S, Tutkur D, Bodenlenz M, Pieber TR. Periodic extraction of interstitial fluid from the site of subcutaneous insulin infusion for the measurement of glucose: a novel single-port technique for the treatment of type 1 diabetes patients. Diabetes Technol Ther 2013 Jan;15(1):50-9.

124.	Nørgaard K, Shin J, Welsh JB, Gjessing H. Performance and acceptability of a combined device for insulin infusion and glucose sensing in the home setting. J Diabetes Sci Technol 2015 Mar;9(2):215-20.

125.	Haidar A, Smaoui MR, Legault L, Rabasa-Lhoret R. The role of glucagon in the artificial pancreas. Lancet Diabetes Endocrinol 2016 Jun;4(6):476-9.

126.	Haidar A, Legault L, Matteau-Pelletier L, Messier V, Dallaire M, Ladouceur M, et al. Outpatient overnight glucose control with dual-hormone artificial pancreas, single-hormone artificial pancreas, or conventional insulin pump therapy in children and adolescents with type 1 diabetes: an open-label, randomised controlled trial. Lancet Diabetes Endocrinol 2015 Aug;3(8):595-604.

127.	Russell SJ, El-Khatib FH, Sinha M, Magyar KL, McKeon K, Goergen LG, et al. Outpatient glycemic control with a bionic pancreas in type 1 diabetes. N Engl J Med 2014 Jul;371(4):313-25.

128.	Jackson MA, Caputo N, Castle JR, David LL, Roberts CT, Ward WK. Stable liquid glucagon formulations for rescue treatment and bi-hormonal closed-loop pancreas. Curr Diab Rep 2012 Dec;12(6):705-10.

129.	Weinzimer SA, Sherr JL, Cengiz E, Kim G, Ruiz JL, Carria L, et al. Effect of pramlintide on prandial glycemic excursions during closed-loop control in adolescents and young adults with type 1 diabetes. Diabetes Care 2012 Oct;35(10):1994-9.

130.	Renukuntla VS, Ramchandani N, Trast J, Cantwell M, Heptulla RA. Role of glucagon-like peptide-1 analogue versus amylin as an adjuvant therapy in type 1 diabetes in a closed loop setting with ePID algorithm. J Diabetes Sci Technol 2014 Sep;8(5):1011-7.

131.	Sherr JL, Patel NS, Michaud CI, Palau-Collazo MM, Van Name MA, Tamborlane WV, et al. Mitigating meal-related glycemic excursions in an insulin-sparing manner during closed-loop insulin delivery: The beneficial effects of adjunctive pramlintide and liraglutide. Diabetes Care 2016 May.

132.	Malek R, Davis SN. Novel methods of insulin replacement: The artificial pancreas and encapsulated islets. Rev Recent Clin Trials 2016;11(2):106-23.

133.	Rege NK, Phillips NFB, Weiss MA. Development of glucose-responsive 'smart' insulin systems. Curr Opin Endocrinol Diabetes Obes 2017 Aug;24(4):267-78.

134.	Valdes-Gonzalez R, Rodriguez-Ventura AL, White DJ, Bracho-Blanchet E, Castillo A, Ramírez-González B, et al. Long-term follow-up of patients with type 1 diabetes transplanted with neonatal pig islets. Clin Exp Immunol 2010 Dec;162(3):537-42.

135.	Elliott RB, Escobar L, Tan PL, Muzina M, Zwain S, Buchanan C. Live encapsulated porcine islets from a type 1 diabetic patient 9.5 yr after xenotransplantation. Xenotransplantation 2007 Mar;14(2):157-61.

136.	Basta G, Montanucci P, Luca G, Boselli C, Noya G, Barbaro B, et al. Long-term metabolic and immunological follow-up of nonimmunosuppressed patients with type 1 diabetes treated with microencapsulated islet allografts: four cases. Diabetes Care 2011 Nov;34(11):2406-9.